## Foglio1
Speculum, vaginal, Graves, (Several sizes) (the specification detailed below is including different models and not only "Graves")
i | Version No. | 1
ii | Date of initial version | December 2019
iii | Date of last modification
iv | Date of publication | February 2020
v | Completed / submitted by | WHO
Name, category and coding
1 | WHO Category / Code | Speculum, vaginal: XD9478 - (Speculum, vaginal, reusable: XD1KF1)
2 | Generic name | Speculum, vaginal
3 | Specific type or variation (optional) | Vaginal speculum, bivalved
4 | GMDN name
5 | GMDN code
6 | GMDN category
7 | UMDNS name
8 | UMDNS code
9 | UNSPS code (optional)
10 | Alternative name/s (optional)
11 | Alternative code/s (optional)
12 | Keywords (optional) | Surgical, gynaecology, vaginal exam
13 | GMDN/UMDNS definition (optional)
Purpose of use
14 | Clinical or other purpose | A vaginal speculum is a device intended to open the vaginal canal to enable a healthcare provider to visually inspect the cervix and collect vaginal or cervical specimens and/or perform surgical operations in a woman’s lower genital tract.
15 | Level of use (if relevant) | Hospital, clinic, or health post
16 | Clinical department/ ward(if relevant) | Family medicine; gynaecology; obstetrics; outpatient clinic; outreach.
17 | Overview of functional requirements | Bivalved vaginal specula have two blades that are self-retaining to facilitate visualization of the cervix for observation, testing, or to carry out a procedure.
Technical characteristics
18 | Detailed requirements | • Specify: reusable or single-use • Bivalved and self-retaining to maintain an open vaginal canal • To be available in a variety of sizes (large, medium, small, AND narrow e.g. Pedersen) • Type to be specified based on user preference or availability Example types and sizes: Collins: o Large – blade length: 110mm (+/- 5%), blade width: 40mm (+/- 5%). o Medium – blade length: 100mm(+/- 5%), blade width: 35mm (+/- 5%). o Small – blade length: 85mm(+/- 5%), blade width: 30mm (+/- 5%). Cusco: o Large – blade length: 11.5cm (+/- 5%), blade width: 3.5cm (+/- 5%). o Medium – blade length: 9.5cm (+/- 5%), blade width: 3.5cm (+/- 5%). o Small – blade length: 7.5cm (+/- 5%), blade width: 2cm (+/- 5%). Graves: o Large – blade length: 115mm (+/- 5%), blade width: 35mm (+/- 5%). o Medium – blade length: 95mm(+/- 5%), blade width: 35mm (+/- 5%). o Small – blade length: 75mm(+/- 5%), blade width: 20mm (+/- 5%). Pederson (narrower version of graves): o Large – blade length: 115mm (+/- 5%), blade width: 25mm (+/- 5%). o Medium – blade length: 95mm(+/- 5%), blade width: 22mm (+/- 5%). o Small – blade length: 75mm(+/- 5%), blade width: 13mm (+/- 5%). * More sizes and types available (e.g. larger, smaller, wider)
19 | Displayed parameters | N.A.
20 | User adjustable settings | N.A.
Physical/chemical characteristics
21 | Components(if relevant) | N.A.
22 | Mobility, portability(if relevant) | N.A.
23 | Raw Materials(if relevant) | Reusable: metal alloys (typically non-quenched, non-magnetic, austenitic stainless steel) and shall be autoclavable. Single-use: highstrength plastic (for example acrylics) and are supplied as sterile.
Utility requirements
24 | Electrical, water and/or gas supply (if relevant) | N.A.
Accessories, consumables, spare parts, other components
25 | Accessories (if relevant) | N.A.
26 | Sterilization process for accessories (if relevant) | N.A.
27 | Consumables / reagents (if relevant) | N.A.
28 | Spare parts (if relevant) | N.A.
29 | Other components (if relevant) | N.A.
Packaging
30 | Sterility status on delivery (if relevant) | Reusable: N/A Single-use: sterile
31 | Shelf life (if relevant) | N.A.
32 | Transportation and storage (if relevant) | Storage area should be clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 85% (non-condensing), and be protected from dripping water.
33 | Labelling (if relevant) | Labelling on the primary packaging should include the name and/or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements.
Environmental requirements
34 | Context-dependent requirements | Environmental conditions vary globally and can be extreme; however, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
Training, installation and utilization
35 | Pre-installation requirements(if relevant) | N.A.
36 | Requirements for commissioning (if relevant) | N.A.
37 | Training of user/s (if relevant) | Clinical staff training in vaginal examinations and gynaecological procedures.
38 | User care(if relevant) | N.A.
Warranty and maintenance
39 | Warranty | One year
40 | Maintenance tasks | N.A.
41 | Type of service contract | N.A.
42 | Spare parts availability post-warranty | N.A.
43 | Software / Hardware upgrade availability | N.A.
Documentation
44 | Documentation requirements | • Instructions for use and service manuals to be provided (including procedures for decontamination). • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local agent.
Decommissioning
45 | Estimated Life Span | N.A.
Safety and standards
46 | Standards, for the manufacturer and the equipment | US FDA: Device Class 1 for metal speculum - US FDA: Device Class 2 for non-metal speculum EU: Class I
47 | Regulatory Approval / Certification | Compliance to (where applicable, but not limited to): • National Regulatory Agency/Authority (NRA) requirements compliance • Approval by regulatory body of country of manufacturer (if applicable) And at least one of: • FDA 510k clearance (US FDA) • CE mark (EU), with indication of Notifying Body (when applicable) • Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan.
48 | International standards (for reference) | Compliant with active version of the following standards (or equivalent): For reusable products: • ISO 7153-1: Surgical instruments – Materials -- Part 1: Stainless steel • ISO 13402: Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. For singe-use products, supplied as sterile: • ISO 17664: Processing of health care products – Information to be provided by the medical manufacturer for the processing of medical devices • ISO 11135: Sterilization of health-care products - Ethylene oxide • ISO 11137: Sterilization of health care products - Radiation • ISO 11607: Packaging for terminally sterilized medical devices
49 | Regional/local standards | Country-specific and regional standards may apply
50 | Regulations | US regulations: 21 CFR part 820 Quality System Regulation 21 CFR part 878.1800 - Speculum and accessories. EU regulations: European Commission Regulation (EU) No. 2017/745 (replacing 93/42/EEC) "