## Foglio1
Airway, oropharyngeal, Guedel (different sizes)
Published on: 19 November 2020 - Version: 1
WHO Item name: | Required technical specifications: | Dimensions / Sizes / Presentation: | Material: | Certificates for product safety and quality assurance, as requested or equivalent
Airway, oropharyngeal, Guedel, adolescent, size 3 (80 mm), autoclavable | One-piece, semi-rigid, curved plastic tube, used to be inserted through the oropharynges to facilitate airway management. Guedel type. Flange surface is permanently marked with tube size/length in mm, and the manufacturer or supplier’s name. It is bite resistant. Proximal (or buccal) end straight and reinforced. Distal end semi-rigid, curved, with atraumatic soft rounded edges. | Size/length /ISO graduation/inner diameter/ colour code/patient: no. 3/80 mm/8/4.5 mm/green/ for small adult. | Latex-free. Polyethylene vinyl acetate (EVA) and polyvinyl chloride (PVC); siliconized; transparent; medical grade. Hygienically clean for medical use. It must resist steam sterilization at 121 °C or 134 °C. | EN12181 or ISO 5364: Oropharyngeal airways. ISO10993-1: Biological safety testing.
Airway, oropharyngeal, Guedel, adult, size 4 (90 mm), autoclavable | One-piece, semi-rigid, curved plastic tube, used to be inserted through the oropharynges to facilitate airway management. Guedel type. Flange surface is permanently marked with tube size/length in mm, and the manufacturer or supplier’s name. It is bite resistant. Proximal (or buccal) end straight and reinforced. Distal end semi-rigid, curved, with atraumatic soft rounded edges. | Size/length/ISO graduation/inner diameter/ colour code/patient: no. 4/90 mm/9/4.5 mm/ yellow/for adult. | Polyethylene vinyl acetate (EVA) and polyvinyl chloride (PVC); siliconized; transparent; medical grade. Hygienically clean for medical use. It must resist steam sterilization at 121 °C or 134 °C. | EN12181 or ISO 5364: Oropharyngeal airways. ISO10993-1: Biological safety testing.
Airway, oropharyngeal, Guedel, adult, size 5 (100 mm), autoclavable | One-piece, semi-rigid, curved plastic tube, used to be inserted through the oropharynges to facilitate airway management. Guedel type. Flange surface is permanently marked with tube size/length in mm, and the manufacturer or supplier’s name. It is bite resistant. Proximal (or buccal) end straight and reinforced. Distal end semi-rigid, curved, with atraumatic soft rounded edges. | Size/length/ISO graduation/inner diameter/ colour code/patient: no. 5/100 mm/10/5.0 mm/red/ for large adult. | Polyethylene vinyl acetate (EVA) and polyvinyl chloride (PVC); siliconized; transparent; medical grade. Hygienically clean for medical use. It must resist steam sterilization at 121 °C or 134 °C. | EN12181 or ISO 5364: Oropharyngeal airways. ISO10993-1: Biological safety testing.
Airway, oropharyngeal, Guedel, child, size 2 (70 mm), autoclavable | One-piece, semi-rigid, curved plastic tube, used to be inserted through the oropharynges to facilitate airway management. Guedel type. Flange surface is permanently marked with tube size/length in mm, and the manufacturer or supplier’s name. It is bite resistant. Proximal (or buccal) end straight and reinforced. Distal end semi-rigid, curved, with atraumatic soft rounded edges. | Size/length/ISO graduation/inner diameter/ colour code/patient: no. 2/70 mm/7/4.0 mm/white/ for young adult. | Polyethylene vinyl acetate (EVA) and polyvinyl chloride (PVC); siliconized; transparent; medical grade. Hygienically clean for medical use. It must resist steam sterilization at 121 °C or 134 °C. | EN12181 or ISO 5364: Oropharyngeal airways. ISO10993-1: Biological safety testing.