## Foglio1
Cardiac Ablation Unit (the technical specification below is a general reference, it is not specific for "cardiac" applications but more focused on obstetric/gynaecologic surgery)
i | Version No. | 1
ii | Date of initial version | December 2019
iii | Date of last modification
iv | Date of publication | February 2020
v | Completed / submitted by | WHO
Name, category and coding
1 | WHO Category / Code
2 | Generic name | Thermal ablation device
3 | Specific type or variation (optional) | Portable, handheld thermal ablation device Benchtop thermal ablation device
4 | GMDN name
5 | GMDN code
6 | GMDN category
7 | UMDNS name
8 | UMDNS code
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Thermal coagulator; thermocoagulator; cold coagulator; Semm coagulator; thermoablation device
11 | Alternative code/s (optional)
12 | Keywords (optional) | Thermal ablation device; cervical precancer; destruction of abnormal human cervical tissue.
13 | GMDN/UMDNS definition (optional)
Purpose of use
14 | Clinical or other purpose | A thermal ablation device is a self-contained, electrically powered medical instrument with a probe tip heated to 100°C designed to destroy tissue of the uterine cervix.
15 | Level of use (if relevant) | Handheld device: health posts, clinics or outreach, hospitals Benchtop device: hospitals or clinics with mains electricity
16 | Clinical department/ ward(if relevant) | Family medicine; gynaecology; outpatient clinic; outreach; oncology; obstetrics; surgery; nursing services
17 | Overview of functional requirements | • Probe tip temperature controlled to reach 100°C • Visual and/or audible cues to ensure working temperature reached • Simple and easy to use, appropriate for all levels of care. • Handheld device has added feature of portability with rechargeable batteries.
Technical characteristics
18 | Detailed requirements | Portable, handheld thermal ablation device OR Benchtop thermal ablation device: • Probe tip temperature controlled to reach 100°C • Visual and/or audible cues to ensure working temperature reached • Simple and easy to use, appropriate for all levels of care. • Handheld device has added feature of portability with rechargeable batteries. • Rated IPX1 (console) and IPX7 (therapy probe and or instrument cable on benchtop)
19 | Displayed parameters | User interface should provide audible and/or visual feedback: • Handheld: light/beep when temperature reached • Benchtop: display to indicate temperature.
20 | User adjustable settings | Benchtop: user can adjust temperature between 60°-120°C
Physical/chemical characteristics
21 | Components(if relevant) | Minimum of 2 probe tips required: o One probe must be flat. o The second probe can be either flat or can have a gentle nipple extrusion not exceeding 5mm (so as to anchor in centre of cervix but not to ablate endocervix). o Probes should not have any sharp edges. o Varying diameters, ranging from 8 mm to 25 mm o Biocompatible, material that will not adhere to cervix o Reusable and thus able to be decontaminated Approximate dimensions of the thermal ablation devices: - Handheld: 5 cm (W), 20 cm (H), 5 cm (D) and weighs <400g - Benchtop: 35 cm (W), 15 cm (H), 10 cm (D) and weighs 3.5 kg Both devices come with power adaptors Recommended extra battery for handheld unit
22 | Mobility, portability(if relevant) | Handheld device: Can be easily carried in a case or backpack, with all accessories, all totalling <2kg
23 | Raw Materials(if relevant) | N.A.
Utility requirements
24 | Electrical, water and/or gas supply (if relevant) | • The unit (either benchtop or handheld when charging) is suggested to be connected to a reliable power source • Handheld, battery operated device to have a minimum cumulative run-time of 1 hour on a single charge. • Electrical source requirements (based on country/setting of use): o Amperage: ______; o Voltage: ______; o Plug type: ______ .
Accessories, consumables, spare parts, other components
25 | Accessories (if relevant) | • Speculum • Light source • Timer
26 | Sterilization process for accessories (if relevant) | Probes should withstand repeated cycles of decontamination,done so according to manufacturers’ instructions. Device handle, charging base, benchtop unit, power supply should be cleaned and decontaminated after each use according to manufacturer’s instructions.
27 | Consumables / reagents (if relevant) | N.A.
28 | Spare parts (if relevant) | • Consider extra probes to account for decontamination cycles and patient scheduling. • A backup battery pack is recommended for handheld devices. Note: probes and spares are not interchangeable between devices of different brands and models and can vary in their design and lifetime.
29 | Other components (if relevant) | N.A.
Packaging
30 | Sterility status on delivery (if relevant) | N.A.
31 | Shelf life (if relevant) | N.A.
32 | Transportation and storage (if relevant) | Storage area should be clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 60% (non-condensing), and be protected from dripping water.
33 | Labelling (if relevant) | Labelling on the primary packaging should include the name and/ or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements.
Environmental requirements
34 | Context-dependent requirements | Environmental conditions vary globally and can be extreme; however, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
Training, installation and utilization
35 | Pre-installation requirements(if relevant) | Handheld: no pre-installation requirements. Benchtop: if necessary, clear instructions/diagrams for assembly/reassembly must be included.
36 | Requirements for commissioning (if relevant) | N.A.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Clinical staff training in cervical precancer lesion treatment guidelines and device use to be provided.
38 | User care(if relevant) | Prior to use: • Inspect for visible damage to device (handle, probes, any connections). • Make sure that no parts are missing or loose. • Make sure that connecting elements between instruments function properly.
Warranty and maintenance
39 | Warranty | Minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided.
40 | Maintenance tasks | Handheld device: no user-serviceable parts. Benchtop device: see manual or contact manufacturer for information
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described (if applicable).
42 | Spare parts availability post-warranty | 8 years minimum, starting from installation/commissioning
43 | Software / Hardware upgrade availability | Costs and types of post-warranty service contract available shall be described (if applicable).
Documentation
44 | Documentation requirements | • Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided (if applicable). • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers.
Decommissioning
45 | Estimated Life Span | Handheld device: 7 years (can vary) Benchtop main unit: approximately 10 years if maintenance and service requirements are met.
Safety and standards
46 | Standards, for the manufacturer and the equipment | US FDA: Device Class 2 EU: Class Iia
47 | Regulatory Approval / Certification | Compliance to (where applicable, but not limited to): - National Regulatory Agency/Authority (NRA) requirements compliance - Approval by regulatory body of country of manufacturer (if applicable) And at least one of: - FDA 510k clearance (US FDA) - CE mark (EU), with indication of Notifying Body (when applicable) - Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan.
48 | International standards (for reference) | Compliant with active version of the following standards (or equivalent): General manufacturing: • ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes • ISO 14971: Medical Devices - Application of Risk Management to Medical Devices • ISO 15223-1: Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements Safety & product standards • IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance • IEC 60601-1-2: Medical electrical equipment - Part 1-2 General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests Probe-specific requirements: • ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process • ISO 10993-5: Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity • ISO 10993-10: Biological evaluation of medical devices Part 10: Tests for irritation and skin sensitization Battery-operated only: • IEC 62133 - Secondary cells and batteries containing alkaline or other non-acid electrolytes – Safety requirements for portable sealed secondary cells o Part 1: Nickel o Part 2: Lithium
49 | Regional/local standards | Country-specific and regional standards may apply
50 | Regulations | US regulations: 21 CFR part 820 Quality System Regulation EU regulations: European Commission Regulation (EU) No. 2017/745 (replacing 93/42/EEC)