## Foglio1
HPV DNA Test (Technical Specifications for HPV NAT IVDs)
i | Version No. | 1
ii | Date of initial version | December 2019
iii | Date of last modification
iv | Date of publication | February 2020
v | Completed / submitted by | WHO
Name, category and coding
1 | WHO Category / Code | Under development
2 | Generic name | HPV NAT to be used at point of care
3 | Specific type or variation (optional) | N.A.
4 | GMDN name
5 | GMDN code
6 | GMDN category
7 | UMDNS name
8 | UMDNS code
9 | UNSPS code (optional)
10 | Alternative name/s (optional)
11 | Alternative code/s (optional)
12 | Keywords (optional)
13 | GMDN/UMDNS definition (optional)
Intended use
14 | Test purpose | For screening. To detect nucleic acid for high-risk HPV genotypes including HPV 16 and HPV 18.
15 | Specific disorder/condition or risk factor of interest | Certain subtypes of HPV as the causative agent of cervical and other anogenital carcinomas.
16 | Testing population | Females over 30 years of age at risk of cervical cancer
17 | Level of the health system | IVDs for use at point of care should be able to be used in settings without laboratory infrastructure, e.g. no or intermittant electricity, no reagent grade water, no specialised specimen collection available, no specialised laboratory staff available. Level I, Level II, Level III, Level IV health facilities.
18 | Intended user | Specimen collection - trained healthcare worker or self test procedure - trained healthcare worker
Performance characteristics
19 | Clinical sensitivity | See TSS-4: IVDs used for the detection of high-risk HPV types in cervical cancer screening for performance requirements
21 | Analytical sensitivity
22 | Clinical specificity
23 | Invalid/unreturnable rate
Operational characteristics
24 | Detection type | Qualitative detection of DNA with PV gene target (L1, E1 ORF).
25 | Genotype detection | Relevant genotypes to be detected are HPV 16, 18, 31, 33, 35, 39, 45, 51, 52, 56, 58, 59 (with 66, 68 acceptable but not preferable) as optional pooled result. Detection of HPV 16 or HPV 16 and 18 is required at a minimum. Individual targeted genotyping as a reflex test or concurrent (preferred) with pooled results; if reporting out individual genotypes, must have the ability to distinguish between genotypes.
26 | Specimen type | Cervical or vaginal, generally using device and specimen transport medium validated and supplied by the manufacturer of IVD. The required specimen volume for testing must be specified.
27 | Specimen throughput | For IVDs used at POC, throughput varies based on product design, but is usually no more than 8 specimens per 8 hour shift per operator.
28 | Time to first results (including hands-on time) | Specimens per run, total run time, hands-on-time and walk-away time must be specified.
29 | Displayed parameters | HPV 16 detected/not detected HPV 18 detected/not detected HPV 16 or 18 detected/not detected Pooled result, if performed
30 | Internal quality control | Essential, amplification control Desirable, positive and negative controls Desirable, sample adequacy control (human housekeeping gene)
31 | Compatibility with external quality control material | Compatibility with external quality control material
32 | Shelf life upon manufacture (months) | Shelf-life must be specified.
33 | Guaranteed shelf life upon delivery | Must have a minimum of 80% of the guaranteed shelf life at the time of delivery, where guaranteed shelf life may be less than shelf life upon manufacture. Note: the guaranteed shelf life is highly dependent on INCOTERMs as stipulated in procurement contracts.
34 | Stability of reagents (temperature and humidity) | Transport (room temperature; 17 to 27 °C) Storage (2 to 30 °C) On-board (2 to 30 °C). Manufacturer must provide duration for on-board stability. In-use (2 to 30 °C)
35 | Stability for specimen collection media | Transport (room temperature; 17 to 27 °C) Storage (2 to 30 °C) In-use (2 to 30 °C)
36 | Stability for controls/calibrators | Transport (room temperature; 17 to 27 °C) Storage (2 to 30 °C) On-board (2 to 30 °C). Manufacturer must provide duration for on-board stability. In-use (2 to 30 °C)
37 | Operating conditions for analysers (temperature and humidity) | Manufacturer must provide required operating conditions for the IVD (temperature and humidity)
Phisical/Chemical characteristics
38 | Components(if relevant) | HPV NAT IVDs will normally have the following essential components: 1. Specimen collection device and transportation media 2. Reagent kit(s) or cartridge 3. Analyser(s) for extraction, amplification and detection, separate or combined.
39 | Footprint (cubic m3) | No more than size of average room.
40 | Weight (kg) and volume | Must be able to be placed on a non-reinforced table or else come with own stand.
Infrastructure Requirements
41 | Electricity | Constant electrical supply required with access to an uninterrupted power supply (UPS), or battery powered.
42 | Water (reagent grade) | Required/not required
43 | Refrigeration or cold chain | No cold chain for IVDs used at POC.
Accessories, consumables, spare parts, other components
44 | Consumables / reagents (if relevant) | HPV NATs come in different designs and therefore requirements are dependent on manufacturer specifications. HPV NATs will normally require the following consumables for each test: 1. Specimen collection device 2. Reagents contained within a test kit or a cartridge
45 | Items required but not provided | Specimen collection devices and transport media Specimen racks Bleach 70% ethanol Paper towel Powder-free gloves Pipettes and plugged pipette tips Waste disposal For provider-collected specimens, an exam table with stirrups, adjustable exam light and vaginal speculum are also required.
46 | Other auxiliary laboratory equipment | Vortex
Labelling
47 | Instructions for use | IFU submitted must relate to regulatory version registered for sale and use in country of supply.
48 | Certificate of analysis | Must be submitted with each consignment of reagents shipped.
Environmental and safety requirements
49 | Hazardous classification | Hazardous goods classification, including material safety data sheet (MSDS)
50 | Disposal requirements | Biohazard receptacles are required for disposal of any biological specimen, including amplified genetic material contained within test kits, cartridges, or collection devices.
Training, installation and utilization
51 | Installation and calibration | Any substantive calibration to be conducted by the supplier (should not be separately charged). Only minimal calibration to be conducted by testing provider.
52 | Training of users | Pre-service and in-service
Warranty and maintenance
53 | Warranty | Minimum 24 months.
54 | Preventive maintenance | Expected minimal maintenance to be conducted by testing provider should be stated. Frequency of servicing based on fixed time periods or based on the number of tests the instrument processes.
55 | Corrective maintenance | Major fixes and/or replacements, the response and resolution times are governed by the service level agreement.
56 | Type of service contract | Must cover labour, repair, spare parts, loaner instrument, shipping and logistics costs, and training
57 | Spare parts availability, post-warranty | To be covered in service level agreement.
Decommissioning
58 | Estimated Life Span | Not less than 2 years
Quality and registration
59 | Global regulatory approvals | Careful consideration should be taken to ensure that the selected products have been assessed to an appropriate stringency level based on the risk classification for HPV NAT IVDs.
60 | Free sale certificate | Provide valid certification for export from country of origin
61 | WHO prequalification status | Desirable. Provide valid WHO PQ Public Report and current WHO list of prequalified IVDs.
62 | International standards | Compliant with active version of the following standards (or equivalent): • ISO 13485: Medical devices—Quality management systems (must provide certification) • ISO 14971: Medical devices—Application of risk management to medical devices • ISO 23640: In vitro diagnostic medical devices—Evaluation of stability of in vitro diagnostic reagents • ISO 18113-1: In vitro diagnostic medical devices -- Information supplied by the manufacturer (labelling) -- Part 1: Terms, definitions and general requirements
63 | National registration | Essential. Provide valid certification.