## Foglio1
IV infusion flow regulator, manual or with dia (The specification detailed is for general reference only and it refers to an "Electronic Drop counter")
Published on: 19 November 2020 - Version: 1
Alternative naming: drops monitor; infusion rate monitor.
Technical Specification | Minimum requirements
General technical requirements | Capable to monitor the flow rate of IV fluids by monitoring the drops through the drip chamber of a standard IV administration set. Equipment provided already calibrated (if necessary) and “ready to use”. Provided with display for data and settings easy visualization, LCD preferable. Equipment capable to work with most common tubing sets and drips (adult and paediatric). Ingress protection not less than IP22. Continuous operation within specification in ambient temperature of at least 5–40 °C, relative humidity of at least 10–85% non-condensing. 90% relative humidity (preferable). Equipment with adjustable flow rate (mL/hr) and/or drops per minute and/or total volume administered. Treatment time adjustable (preferable).
Monitored and controlled parameters (by user) Displayed parameters (colour and graphic preferable) | At least the following parameters can be displayed: • calculated flow rate • fluid volume • drops per minute • treatment time (preferable).
Alarms | At least the following video and/or audio alarms available: • battery low • flow rate and/or volume and/or drops per minute set values has been reached • set rate changes (preferable) • empty bag/bottle (preferable).
Consumables, labelled “single use” (included and mentioned in a disaggregated list) | Intravenous compatible tubing set for both adult and paediatric: quantity to be defined.
Accessories, reusable (included and mentioned in a disaggregated list) | As per manufacturer proposal.
Spare parts (included and mentioned in a disaggregated list) | As per manufacturer proposal. Include calibration software and hardware, if applicable Include list of spare parts with their part numbers and costs.
Portability | N/A
Power supply (voltage, frequency and plug vary across the countries) | Operates with built-in rechargeable battery and/or AC mains power: 100–240 V~/50–60 Hz. Automatic switch from AC mains power mode to battery operating mode and vice versa (if applicable). Internal rechargeable battery with at least 24 hours working time. Battery charger provided. Built-in battery charger (preferable). Low battery alarm available.
Documentation (included) | Instruction for use; service manual and product information to be provided in English, at least.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]). National local regulatory approval (of recipient country, as applicable)
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: General medical equipment: IEC 60601-1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance or ANSI/AAMI ES60601-1:2005+A2 (R2012) Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-1:2000 Medical electrical equipment – Part 1-1: General requirements for safety – Collateral standard: Safety requirements for medical electrical systems. IEC 60601-1-2:2007 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests.
Warranty | Minimum 2 years. Availability of accessories, consumables and spare parts for at least 5 years.