## Note
Format of WHO technical specifications
MEDICAL DEVICE SPECIFICATION (including information on the following where relevant/appropriate, but not limited to)
Instructions and examples
i | Version No. | Rows i - v and 1 are completed and managed by WHO
ii | Date of initial version
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | Usually name of Institution / Organization / UN agency
NAME, CATEGORY AND CODING
1 | WHO Category / Code
2 | Generic name | Name of the medical device as commonly used (e.g. anaesthesia machine).
3 | Specific type or variation (optional) | Characteristics of the device that distinguish it from other similar devices or devices of the same generic name (e.g. handheld, bench-top, portable, digital, adult/paediatric/neonatal, consumable/disposable, single-use, etc.).
4 | GMDN name | Name as produced and maintained by the Global Medical Devices Nomenclature (GMDN) Agency, e.g. Anaesthesia unit, mobile. (NB: Access to GMDN Agency nomenclature system may be restricted - see http://www.gmdnagency.com/ for further information).
5 | GMDN code | Comments as for [9]; GMDN code for 'Anaesthesia unit, mobile' is 47769 (all GMDN device codes have 5 digits)
6 | GMDN category | Comments as for [9]; GMDN category for 'Anaesthesia unit, mobile' is '02 Anaesthetic and respiratory devices'.
7 | UMDNS name | Name as produced and maintained by the ECRI institute, e.g. Anaesthesia Units (NB: Access to ECRI nomenclature system may be restricted - see https://www.ecri.org/Products/Pages/UMDNS.aspx for further information).
8 | UMDNS code | Comments as for [12]; ECRI code for 'Anaesthesia Units' is 10134 (all ECRI device codes have 5 digits).
9 | UNSPS code (optional) | United Nations Standard Products and Services Code [ see http://www.unspsc.org/ ]. This coding system uses a hierarchy of Family-Class-Commodity. For an anaesthesia unit, which comprises a number of functional modules, there are a number of corresponding Commodity codes and titles listed under more than Class; e.g. Commodities 42272501 'Gas anaesthesia apparatus' and 42272502 'Absorber units for gas anaesthesia apparatus' are included under Class 42272500 'Anaesthesia apparatus and accessories and supplies' in the Family 42270000 'Respiratory and anaesthesia and resuscitation products'.
10 | Alternative name/s (optional) | Name/s set by a regional or national authority, local names (e.g. Boyle's machine) or synonyms of formal nomenclature (e.g. anaesthesia apparatus or system).
11 | Alternative code/s (optional) | Corresponding code/s set by a regional or national authority.
12 | Keywords (optional) | Specific area / disease related to the device (e.g. anaesthesia, intra-operative care, etc.).
13 | GMDN/UMDNS definition (optional) | Definitions produced and maintained by the GMDN Agency and ECRI Institute, respectively.
PURPOSE OF USE
14 | Clinical or other purpose | A description of the essential clinical or other objective/s associated with the device's utilisation, e.g. anaesthesia units (allow the anaesthetist to) dispense a mixture of gases and vapours and vary the proportions thereof to control a patient’s level of consciousness and/or analgesia during surgical procedures.
15 | Level of use (if relevant) | The level of healthcare service delivery at which the device is to be used, or is typically used. [ NB: Since the level of skill/s required of the device user/s should also be considered, and the levels of service delivery are not globally standardised, this level may vary from country to country.] Home use should also be considered as a level of care. For our example, the anaesthesia unit would typically be used at district, regional and tertiary hospitals.
16 | Clinical department/ward(if relevant) | The usual service area / functional department in which the device would be used, e.g. Operating room, Intensive Care Unit, Paediatric ward, Outpatient department). Home use should also be considered as a level of care.
17 | Overview of functional requirements | General description of the device's function, e.g. for anaesthesia unit this would include gas/vapour delivery platform; ventilator with patient breathing circuit; scavenging system to capture and exhaust waste gases; physiological and multi-gas monitors, etc.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | The required characteristics and specific/critical functional requirements. e.g. modules, components, measured and/or delivered parameters and associated values and ranges, compatibility / inter-operability requirements, etc.
19 | Displayed parameters | User interface information requirements (e.g. display of pressure, volume, flow, status indicators, inspiration and expiration times, etc.) and format (continuous waveform display, digital, trends, etc.).
20 | User adjustable settings | Device functional parameters, alarms, language, etc. that should be adjustable at the discretion of the user/s.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Dimensions, configuration of complex equipment, etc.
22 | Mobility, portability(if relevant) | Requirements for non-fixed/installed devices, e.g. weight, handles, on castor wheels of specified diameter, etc.
23 | Raw Materials(if relevant) | Applies mainly to surgical instruments and /or implants. e.g. stainless steel (linked to bio-compatibility/patient safety, corrosion resistance, etc.).
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical supply: e.g. nominal mains voltage with frequency and permitted fluctuations, battery operation (if relevant); Water and gas supply: quality and flow rate requirements.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Accessories needed (type, number, functional requirements, etc.) for full and proper functioning of the device.
26 | Sterilization process for accessories (if relevant) | Preferred method to be specified, if appropriate; otherwise to be clearly indicated by manufacturer/supplier.
27 | Consumables / reagents (if relevant) | Consumables (renewables) and disposables (including single-use accessories) to be used with the medical device. Where appropriate, quantity required, shelf life, etc. should be specified.
28 | Spare parts (if relevant) | It would be very useful to know what parts are likely to be needed in the first year of operation (based on average usage and experience elsewhere) and/or in the year after expiry of the warranty period.
29 | Other components (if relevant) | Complementary equipment (e.g. printers, stands, wall mounts, etc.).
PACKAGING
30 | Sterility status on delivery (if relevant) | To be specified - applies to implantables or single-use devices that are delivered sterile
31 | Shelf life (if relevant) | Shelf life and number of uses of the device to be specified
32 | Transportation and storage (if relevant) | Specific considerations for transportation and storage
33 | Labelling (if relevant) | Specific labelling requirements
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Storage and operating temperatures (specify ranges), resistance to high humidity and/or dust levels (specify requirements) - in accordance with local/anticipated conditions.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Construction / structural changes, utility requirements, etc.
36 | Requirements for commissioning (if relevant) | Manufacturer/supplier to perform installation, safety and operation checks before handover. Acceptance tests to be specified and local clinical and technical staff to verify proper and full functioning of device.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Training of maintenance personnel (if relevant) also to be specified and provided.
38 | User care(if relevant) | Information to be provided by manufacturer/supplier, e.g. cleaning, disinfection/sterilization method (for reusable devices).
WARRANTY AND MAINTENANCE
39 | Warranty | Date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
40 | Maintenance tasks | Specific equipment for needed for calibration or testing purposes must be specified. Advanced maintenance tasks required shall be documented, with details of maintenance support from manufacturer/supplier.
41 | Type of service contract | Non Comprehensive or Comprehensive Contract
42 | Spare parts availability post-warranty | Usually at least 5 years after device acquisition.
43 | Software / Hardware upgrade availability | To be specified.
DOCUMENTATION
44 | Documentation requirements | Operating and service manuals (language/s to be specified) including lists of important spares and accessories - with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
DECOMMISSIONING
45 | Estimated Life Span | Predictable average life span, if it is assumed the average frequency of utilization, maintenance and failure. The device would be better to be assessed on the replacement concerning this span. (see 'How to Plan and Budget for your Healthcare Technology' http://www.who.int/management/plan_budget_healthcare.pdf)
SAFETY AND STANDARDS
46 | Standards, for the manufacturer and the equipment | Manufaturer standards about quality issues such as ISO 9001, and ISO 13485, etc.
47 | Regulatory Approval / Certification | e.g. FDA approval (USA), CE mark (EU)
48 | International standards | Specified for compliance by manufacturers in global marketplace, notably ISO 13485: Quality Management System and ISO 14971: Risk Management System. Apply to categories of devices, e.g. for electromedical devices IEC 60601-1 (General requirements for basic safety and essential performance), IEC 60601-1-1-1 (Collateral standard: safety requirements for medical electrical systems) and IEC 60601-1-2 (Collateral standard: Electromagnetic compatibility - Requirements and tests). Apply to specific devices, e.g. IEC 60601-2-19 (Particular requirements for the basic safety and essential performance of infant incubators), ISO 10079-1 (Medical suction equipment), etc.

## Defibrillator
DEFIBRILLATOR, EXTERNAL, MANUAL OR SEMI-AUTOMATED
i | Version No. | 2
ii | Date of initial version | 13/06/2012 (modified on 04/07/2014)
iii | Date of last modification | 2021-10-06
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Defibrillator
3 | Specific type or variation (optional) | External, manual, semi-automated
4 | GMDN name | Rechargeable professional automated external defi brillator
5 | GMDN code | 48048
6 | GMDN category | In vitro diagnostic devices
7 | UMDNS name | Defibrillator/Cardioverters, Defibrillator/Pacemakers, External, Defibrillators, External, Manual
8 | UMDNS code | 18-502, 17-882, 11-134
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Battery-powered defibrillators, cardioverters, defibrillator/monitor/pacemakers, external biphasic defibrillators, external monophasic defibrillators, monitor/defibrillators, AED (automatic external defibrillator); AED
11 | Alternative code/s (optional) | MS 37600; MS 43180; MS 42557; MS 17116; MS 34110
12 | Keywords (optional) | Defibrillation, AED, cardioverter, pacemaker
13 | GMDN/UMDNS definition (optional) | A portable electronic device designed to automatically detect cardiac arrhythmias (ventricular fibrillation / pulseless ventricular tachycardia) in a sudden cardiac arrest patient, and to audibly / visually instruct an operator to enable it to activate defibrillation of the heart or allow the operator to decide when to activate defibrillation based on its electrocardiogram (ECG) display. It is intended to be used by healthcare professionals (e.g., paramedics, medical staff) in healthcare settings. It consists of an external pulse generator and skin-adhesive electrodes to monitor the rhythm/deliver the shocks; it has internal non-rechargeable batteries that must be replaced when indicated.
14 | EMDN code | Z120305, Z12030503
15 | EMDN nomenclature | DEFIBRILLATORS, AUTOMATIC DEFIBRILLATORS
PURPOSE OF USE
16 | Clinical or other purpose | Designed to automatically detect cardiac arrhythmias in a sudden cardiac arrest patient, and to audibly / visually instruct an operator to enable it to activate defibrillation of the heart or allow the operator to decide when to activate defibrillation based on its electrocardiogram (ECG) display.
17 | Level of use (if relevant) | Health centers, district hospital, provincial hospital, specialized hospital.
18 | Clinical department/ward(if relevant) | Outpatient, Emergency (ER), Operating theaters, Intensive Care Units (ICU)
19 | Overview of functional requirements | ECG monitored through either separate pads or handheld defibrillation paddles. Displays and analyzes the ECG and advises operator of patient state.Charges in preparation for shock delivery, after which the operator activates the discharge.Discharge is through handheld paddles connected by extendable wires to the unit.Capable of discharging either directly or synchronized with ECG. It can be operated off mains or internal battery power. Manual override function included. Printing functionalities. Additional monitoring capabilities can be included.
TECHNICAL CHARACTERISTICS
20 | Detailed requirements | Manual and semi-automatic operating modes. Selft-Test and discharge test modes available Biphasic waveform operation. Pediatric attenuating system available Maximum energy to be at least 220 joules for adult and 50 joules for pediatric. Conductive area for paddles (in case of manual defibrillator) shall be >50 cm2 for adutl and >15 cm2 for paediatric Electrodes (in case of AED) shall be > 75 cm2 for adults and > 40 cm2 for pediatric. ECG analysis time to be <15 sec. Charge time to full energy to be <10 sec. More than 100 full energy discharges to be possible solely off battery operation (fully charged). Voice prompting function included for operator direction Number of discharges (total lifetime and on current battery) to be displayed. Self test facility to be included. Automatic impedance compensation. External defibrillation discharging start control just only by pressing both buttons on the external paddles. One set of reusable adult external paddles and related pediatric adapters compatible with the equipment. "External (Non-Invasive) pacing capability with at least the following characteristics (approximate values): - Pulse Width: 20 ms, ±5% - Pacing Mode: Demand or fixed - Pacing rate: 40 ppm to 170 ppm - Pacing output: 0 mA to 200 mA, ±5% or 5 mA, whichever is greater. Monitor with large high resolution LCD display for easy viewing of ECG and vital signs. At least the following parameters monitored: ECG, Heart rate, SPO2 pulse oxymetry, NIBP (preferable), Respiration rate (preferable), EtCO2 (preferable). With Automatic External Defibrillator(AED) function. Data storage capability: able to store the last 100 events with date and time. Integrated thermal printer with printing speed at least 25 and 50 mm/s.
21 | Displayed parameters | Indicator for power and battery state required. Additional monitored parameters (i.e. ECG, vital signs) shall be displayed (when available)
22 | User adjustable settings | Energy level delievered in Joules
PHYSICAL/CHEMICAL CHARACTERISTICS
23 | Components(if relevant) | Equipment weight not greater than 10 kg
24 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand. Crash resistant, dust-proof and washable carry-case for the complete equipment provided.
25 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
26 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: 100-240 V, 50/60 Hz. Appropriate external device to protect the equipment against over-voltage and over-current line conditions (between plug and socket). Main cable to be at least 3 m length. Battery compartment to be clearly labeled, splash-proof and requiring tool for opening for battery replacement. Rechargeable batteries with at least the following characteristics: - Automatic switch from electric-line mode to battery operating mode - Numer of discharges with maximal energy in battery operating mode not less than 15 - Continuos monitoring working time in battery operating mode not less than 2 hours - Integrated batteries charger - Low battery visual alarm - 100% high capacity batteries with re-charging time no greater than 10 hours.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
27 | Accessories (if relevant) | Adult and paediatric defibrillation paddles (for manual defibrillator) and electrodes, when applicable (for AED). ECG cables sets for adult and pediatric. All necessary monitoring supplies including ECG cables and wires, NIBP cuffs, SpO2 sensors, etc. (when applicable). Set of spare batteries. External batteries charging station (when applicable) Carry case
28 | Sterilization process for accessories (if relevant) | Not required
29 | Consumables / reagents (if relevant) | Tubes or bottles of coupling gel. Sets ECG pads for adult and pediatric. Electrodes for AED adult and pediatric (when applicable) Thermal printing paper rolls, when applicable.
30 | Spare parts (if relevant) | 1-year spare parts kit as per preventive maintenance programme
31 | Other components (if relevant) | N/A
PACKAGING
32 | Sterility status on delivery (if relevant) | N/A
33 | Shelf life (if relevant) | N/A
34 | Transportation and storage (if relevant) | Unit shall be supplied protectively packed for safe transportation and delivery.
35 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
36 | Context-dependent requirements | Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90% Capable of being stored in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%
TRAINING, INSTALLATION AND UTILISATION
37 | Pre-installation requirements(if relevant) | Not required
38 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation Supplier to perform installation, safety and operation checks before handover
39 | Training of user/s (if relevant) | Training of users in operation and technicians in basic maintenance shall be provided
40 | User care(if relevant) | The whole unit is to be cleanable with alcohol or chlorine wipes or with any standard hospital disinfection procedure/material
WARRANTY AND MAINTENANCE
41 | Warranty | Not less than 2 years. Specific inclusions and exclusions to be listed.
42 | Maintenance tasks | List of equipment and procedures required for local routine maintenance should be provided.
43 | Type of service contract | Costs and types of post-warranty service contract available should be described (when needed). Contact details of manufacturer, supplier and local service agent to be provided.
44 | Spare parts availability post-warranty | Availability of accesories, consumables and spare parts for at least 5 years. Guaranteed time period of availability of spare parts post-warranty should be pointed out.
45 | Software / Hardware upgrade availability | Provided, when necessary
DOCUMENTATION
46 | Documentation requirements | User / technical manual to be supplied in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration, if necessary, and routine maintenance. Battery disposal instructions according local laws.
DECOMMISSIONING
47 | Estimated Life Span | Not less than 5 years
SAFETY AND STANDARDS
48 | Standards, for the manufacturer and the equipment | NMPA certificate provided (Chinese Suppliers only). Certified Quality Management System for medical devices (e.g. ISO 13485, or Good Manufacturing Practice - GMP). Application of risk management to medical devices (e.g. ISO 14971). General quality management (e.g. ISO 9001: Quality management systems – Requirements).
49 | Regulatory Approval / Certification | Free Sales Certificate (FSC) provided by any of the following countries: Australia, Canada, Japan, USA and European Community (e.g. FDA and/or CE certificate given by a third certified party for the specific medical devices proposed (no only declaration of conformity).
50 | International standards | Compliance to the following international standards, when applicable, or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended but compliance to previous standards versions could be considered. IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-1:2015 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems. IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. IEC 60601-2-4 Ed. 3.0:2010 8b) Medical electrical equipment - Part 2-4: Particular requirements for the basic safety and essential performance of cardiac defibrillators