## Foglio1
Infrared Thermometers
i | Version No. | 2
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2020-07-15
iv | Date of publication | 2020-11-19
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Thermometer, infrared, skin
3 | Specific type or variation (optional) | Skin
4 | CND code (https://ec.europa.eu/health/md_topics-interest/overview_en) | V03010102
5 | CND nomenclature | ELECTRONIC THERMOMETERS AND END CAPS
Purpose of use
6 | Clinical or other purpose | Estimate the temperature of a site on the skin.
7 | Level of use (if relevant) | Health post, health centre, district hospital, provincial hospital, specialized hospital, outreach (mobile clinics)
8 | Clinical department/ ward(if relevant) | Emergency room (ER), neonatal intensive care unit (NICU), surgery, outpatient, intensive care unit (ICU), hospital triage and other departments
9 | Overview of functional requirements | Displays patient temperature by measurement of infrared radiation from the skin. Device must be reusable, with sterilizable surface. Display should be easily readable in all levels of ambient light.
Technical characteristics
10 | Detailed requirements | Specified accuracy to be not higher than 0.2–0.3 °C. Measurement range at least from 30–43 °C. High/low patient temperature display feature preferred. Auto power off required after minimum of 1 minute. Out of range indication required. Response (measurement) time not higher than 3 sec. Ready-to-use after switch-on in a time not higher than 10 sec. Infrared (IR) spectral response 6000–14 000 nm. Optimal measuring distance approximately 8–12 cm/4–6 inch. Equipment factory calibrated and pre-set emissivity data for all skin types. Automatic self-test on switch-on. Video and/or audio alert/signal at least for the following cases: switch-on, ready-to-use and measurement completed.
11 | Displayed parameters | Display graded in 0.1/0.3 °C steps. High/low patient temperature. Low battery. Malfunction. °F or °C measurement units.
12 | User adjustable settings | None.
Physical/chemical characteristics
13 | Components(if relevant) | Supplied in protective case for clean storage and safe transport. Unit case should be hard and splashproof. Must be lightweight and comfortable to hold. There must be no sharp edges on the unit.
14 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand.
15 | Raw Materials(if relevant) | N/A
Utility requirements
16 | Electrical, water and/or gas supply (if relevant) | Powered by internal, rechargeable, replaceable battery. Battery cover to be secure but simple to open. Battery to allow at least 4000 measurements between charges. Battery charger to operate from input supply 110–220 V, 60–50 Hz, ± 10% (battery charger built-in or external).
Accessories, consumables, spare parts, other components
17 | Accessories (if relevant) | Full range of any adaptors required to allow for measurement of all ages of patient.
18 | Sterilization process for accessories (if relevant) | Not required.
19 | Consumables / reagents (if relevant) | Not required.
20 | Spare parts (if relevant) | Replacement battery pack, supplied empty of charge.
21 | Other components (if relevant) | N/A
Packaging
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | N/A
24 | Transportation and storage (if relevant) | Unit shall be supplied protectively packed for safe transportation and delivery.
25 | Labelling (if relevant) | N/A
Environmental requirements
26 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0–50 °C and relative humidity of 15–85%, preferably 90%. Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85%, preferably 90%.
Training, installation and utilization
27 | Pre-installation requirements(if relevant) | Not required.
28 | Requirements for commissioning (if relevant) | Safety and operation checks before handover.
29 | Training of user/s (if relevant) | Training of users in operation and technicians in basic maintenance.
30 | User care(if relevant) | The whole unit is to be cleanable with alcohol or chlorine wipes or with any standard hospital disinfection procedure/material.
Warranty and maintenance
31 | Warranty | Not less than 2 years. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided.
32 | Maintenance tasks | List of procedures required for local routine maintenance should be provided.
33 | Type of service contract | Costs and types of post-warranty service contract available should be described (when needed).
34 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty should be pointed out.
35 | Software / Hardware upgrade availability | Not required.
Documentation
36 | Documentation requirements | User/technical manual to be supplied in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration, if necessary, and routine maintenance. Battery disposal according local laws.
Decommissioning
37 | Estimated Life Span | Not less than 5 years.
Safety and standards
38 | Standards, for the manufacturer and the equipment | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (e.g. ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 Medical devices – Application of risk management to medical devices).
39 | Regulatory Approval / Certification | Free sales certificate (FSC). Certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
40 | International standards | Compliance to the following international standards, when applicable, or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended but compliance to previous standards versions could be considered. IEC 60601-1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance IEC 60601-1-2:2007 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests ISO 80601-2-56:2009 Medical electrical equipment – Part 2-56: Particular requirements for basic safety and essential performance of clinical thermometers for body temperature measurement IEC 80601-2-59 Ed. 1.0:2008 (b) Medical electrical equipment – Part 2-59: Particular requirements for the basic safety and essential performance of screening thermographs for human febrile EN ISO 15223-1 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements ASTM E1104-98(2016) Standard Specification for Clinical Thermometer Probe Covers and Sheaths. ASTM E1112-00(2018) Standard Specification for Electronic Thermometer for Intermittent Determination of Patient Temperature.