## Note
Format of WHO technical specifications
MEDICAL DEVICE SPECIFICATION (including information on the following where relevant/appropriate, but not limited to)
Instructions and examples
i | Version No. | Rows i - v and 1 are completed and managed by WHO
ii | Date of initial version
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | Usually name of Institution / Organization / UN agency
NAME, CATEGORY AND CODING
1 | WHO Category / Code
2 | Generic name | Name of the medical device as commonly used (e.g. anaesthesia machine).
3 | Specific type or variation (optional) | Characteristics of the device that distinguish it from other similar devices or devices of the same generic name (e.g. handheld, bench-top, portable, digital, adult/paediatric/neonatal, consumable/disposable, single-use, etc.).
4 | GMDN name | Name as produced and maintained by the Global Medical Devices Nomenclature (GMDN) Agency, e.g. Anaesthesia unit, mobile. (NB: Access to GMDN Agency nomenclature system may be restricted - see http://www.gmdnagency.com/ for further information).
5 | GMDN code | Comments as for [9]; GMDN code for 'Anaesthesia unit, mobile' is 47769 (all GMDN device codes have 5 digits)
6 | GMDN category | Comments as for [9]; GMDN category for 'Anaesthesia unit, mobile' is '02 Anaesthetic and respiratory devices'.
7 | UMDNS name | Name as produced and maintained by the ECRI institute, e.g. Anaesthesia Units (NB: Access to ECRI nomenclature system may be restricted - see https://www.ecri.org/Products/Pages/UMDNS.aspx for further information).
8 | UMDNS code | Comments as for [12]; ECRI code for 'Anaesthesia Units' is 10134 (all ECRI device codes have 5 digits).
9 | UNSPS code (optional) | United Nations Standard Products and Services Code [ see http://www.unspsc.org/ ]. This coding system uses a hierarchy of Family-Class-Commodity. For an anaesthesia unit, which comprises a number of functional modules, there are a number of corresponding Commodity codes and titles listed under more than Class; e.g. Commodities 42272501 'Gas anaesthesia apparatus' and 42272502 'Absorber units for gas anaesthesia apparatus' are included under Class 42272500 'Anaesthesia apparatus and accessories and supplies' in the Family 42270000 'Respiratory and anaesthesia and resuscitation products'.
10 | Alternative name/s (optional) | Name/s set by a regional or national authority, local names (e.g. Boyle's machine) or synonyms of formal nomenclature (e.g. anaesthesia apparatus or system).
11 | Alternative code/s (optional) | Corresponding code/s set by a regional or national authority.
12 | Keywords (optional) | Specific area / disease related to the device (e.g. anaesthesia, intra-operative care, etc.).
13 | GMDN/UMDNS definition (optional) | Definitions produced and maintained by the GMDN Agency and ECRI Institute, respectively.
PURPOSE OF USE
14 | Clinical or other purpose | A description of the essential clinical or other objective/s associated with the device's utilisation, e.g. anaesthesia units (allow the anaesthetist to) dispense a mixture of gases and vapours and vary the proportions thereof to control a patient’s level of consciousness and/or analgesia during surgical procedures.
15 | Level of use (if relevant) | The level of healthcare service delivery at which the device is to be used, or is typically used. [ NB: Since the level of skill/s required of the device user/s should also be considered, and the levels of service delivery are not globally standardised, this level may vary from country to country.] Home use should also be considered as a level of care. For our example, the anaesthesia unit would typically be used at district, regional and tertiary hospitals.
16 | Clinical department/ward(if relevant) | The usual service area / functional department in which the device would be used, e.g. Operating room, Intensive Care Unit, Paediatric ward, Outpatient department). Home use should also be considered as a level of care.
17 | Overview of functional requirements | General description of the device's function, e.g. for anaesthesia unit this would include gas/vapour delivery platform; ventilator with patient breathing circuit; scavenging system to capture and exhaust waste gases; physiological and multi-gas monitors, etc.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | The required characteristics and specific/critical functional requirements. e.g. modules, components, measured and/or delivered parameters and associated values and ranges, compatibility / inter-operability requirements, etc.
19 | Displayed parameters | User interface information requirements (e.g. display of pressure, volume, flow, status indicators, inspiration and expiration times, etc.) and format (continuous waveform display, digital, trends, etc.).
20 | User adjustable settings | Device functional parameters, alarms, language, etc. that should be adjustable at the discretion of the user/s.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Dimensions, configuration of complex equipment, etc.
22 | Mobility, portability(if relevant) | Requirements for non-fixed/installed devices, e.g. weight, handles, on castor wheels of specified diameter, etc.
23 | Raw Materials(if relevant) | Applies mainly to surgical instruments and /or implants. e.g. stainless steel (linked to bio-compatibility/patient safety, corrosion resistance, etc.).
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical supply: e.g. nominal mains voltage with frequency and permitted fluctuations, battery operation (if relevant); Water and gas supply: quality and flow rate requirements.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Accessories needed (type, number, functional requirements, etc.) for full and proper functioning of the device.
26 | Sterilization process for accessories (if relevant) | Preferred method to be specified, if appropriate; otherwise to be clearly indicated by manufacturer/supplier.
27 | Consumables / reagents (if relevant) | Consumables (renewables) and disposables (including single-use accessories) to be used with the medical device. Where appropriate, quantity required, shelf life, etc. should be specified.
28 | Spare parts (if relevant) | It would be very useful to know what parts are likely to be needed in the first year of operation (based on average usage and experience elsewhere) and/or in the year after expiry of the warranty period.
29 | Other components (if relevant) | Complementary equipment (e.g. printers, stands, wall mounts, etc.).
PACKAGING
30 | Sterility status on delivery (if relevant) | To be specified - applies to implantables or single-use devices that are delivered sterile
31 | Shelf life (if relevant) | Shelf life and number of uses of the device to be specified
32 | Transportation and storage (if relevant) | Specific considerations for transportation and storage
33 | Labelling (if relevant) | Specific labelling requirements
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Storage and operating temperatures (specify ranges), resistance to high humidity and/or dust levels (specify requirements) - in accordance with local/anticipated conditions.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Construction / structural changes, utility requirements, etc.
36 | Requirements for commissioning (if relevant) | Manufacturer/supplier to perform installation, safety and operation checks before handover. Acceptance tests to be specified and local clinical and technical staff to verify proper and full functioning of device.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Training of maintenance personnel (if relevant) also to be specified and provided.
38 | User care(if relevant) | Information to be provided by manufacturer/supplier, e.g. cleaning, disinfection/sterilization method (for reusable devices).
WARRANTY AND MAINTENANCE
39 | Warranty | Date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
40 | Maintenance tasks | Specific equipment for needed for calibration or testing purposes must be specified. Advanced maintenance tasks required shall be documented, with details of maintenance support from manufacturer/supplier.
41 | Type of service contract | Non Comprehensive or Comprehensive Contract
42 | Spare parts availability post-warranty | Usually at least 5 years after device acquisition.
43 | Software / Hardware upgrade availability | To be specified.
DOCUMENTATION
44 | Documentation requirements | Operating and service manuals (language/s to be specified) including lists of important spares and accessories - with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
DECOMMISSIONING
45 | Estimated Life Span | Predictable average life span, if it is assumed the average frequency of utilization, maintenance and failure. The device would be better to be assessed on the replacement concerning this span. (see 'How to Plan and Budget for your Healthcare Technology' http://www.who.int/management/plan_budget_healthcare.pdf)
SAFETY AND STANDARDS
46 | Standards, for the manufacturer and the equipment | Manufaturer standards about quality issues such as ISO 9001, and ISO 13485, etc.
47 | Regulatory Approval / Certification | e.g. FDA approval (USA), CE mark (EU)
48 | International standards | Specified for compliance by manufacturers in global marketplace, notably ISO 13485: Quality Management System and ISO 14971: Risk Management System. Apply to categories of devices, e.g. for electromedical devices IEC 60601-1 (General requirements for basic safety and essential performance), IEC 60601-1-1-1 (Collateral standard: safety requirements for medical electrical systems) and IEC 60601-1-2 (Collateral standard: Electromagnetic compatibility - Requirements and tests). Apply to specific devices, e.g. IEC 60601-2-19 (Particular requirements for the basic safety and essential performance of infant incubators), ISO 10079-1 (Medical suction equipment), etc.

## ESU
ELECTROSURGICAL UNIT
i | Version No. | 2
ii | Date of initial version | 13/06/2012 (modified on 04/07/2014)
iii | Date of last modification | 2021-10-06
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Electrosurgical Unit
3 | Specific type or variation (optional) | Bipolar unit; monopolar unit; monopolar/bipolar unit
4 | GMDN name | Electrosurgical system, general-purpose
5 | GMDN code | 44776
6 | GMDN category | Electro mechanical medical devices
7 | UMDNS name | Electrosurgical Units, Monopolar/Bipolar; Electrosurgical Units
8 | UMDNS code | 11490; 18231
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Bovies; Coagulators, Electrosurgical; Diathermy Units, Surgical; Electrocautery Units; Electrosurgical Generators; Endometrial Ablation Sys tems; ESUs; Hyfrecators; Malis Coagulators; Surgical Diathermy Units; Surgical Units; Wapplers; Apparatus, electrosurgical; Surgical diathermy generator.
11 | Alternative code/s (optional) | 531.328.0181 (MX); 531.328.0116 (MX); MS 42361
12 | Keywords (optional) | RF; Coagulator; Electrosurgical; Diathermy; Electrocautery; ESU.
13 | GMDN/UMDNS definition (optional) | An assembly of devices that uses high frequency electrical energy in a radio-frequency (RF) band to develop heat directly within soft-tissue cells (thermodynamic) for cutting and coagulating tissue typically during general surgical procedures. Tissue resistance to the electrical current creates the heat as the current travels through the body between electrodes. The assembly typically includes an energy-producing generator with monitoring functions, a single-use/reusable handpiece with electrodes to apply the energy to the surgical site, connecting cables, and a foot-switch as an option to regulate the energy.
14 | EMDN code | Z12010902
15 | EMDN nomenclature | ELECTROSURGICAL UNITS (ESU) FOR GENERAL USE
PURPOSE OF USE
16 | Clinical or other purpose | Devices intended for surgical cutting and for controlling bleeding by causing coagulation (hemostasis) at the surgical site. Electrosurgery is commonly used in dermatological, gynecological, cardiac, plastic, ocular, spine, ENT, maxillofacial, orthopedic, urological, neuro- and general surgical procedures as well as certain dental procedures.
17 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
18 | Clinical department/ward(if relevant) | Operating room, Emergency room
19 | Overview of functional requirements | Devices intended for surgical cutting (wire electrodes) and for controlling bleeding (ball electrode) by causing coagulation (haemostasis) at the surgical site with the safe passage of electricity at a high frequency to and through tissue. Allows both monopolar and bipolar operation. Adjustable power level can be set by the operator .
TECHNICAL CHARACTERISTICS
20 | Detailed requirements | Radiofrequency generator to be within the range from 300,000 Hz to 5,000,000 Hz, output to be electrically isolated from ground. At least the following modes of operation: - Monopolar - Pure cut: to be at least 300 W - Pure coagulation: to be at least 200 W - Blended: to be at least 200 W - Bipolar: to be at least 50 W - Pinpoint: to be at least 120 W - Spray: to be at least 80 W - Self test mode Equipment provided with at least: 1 monopolar handpiece port, 1 monopolar return electrode port (with alarm when poor contact quality), 1 bipolar outlet and monopolar and bipolar pedal. Operation to be controlled by foot pedal, with minimum 2 meters connection cable, and also by handswitch on probe. Control panel for adjusting and displaying power settings. Visual and audible activation indicators required. Visual and audible cable disconnection alarm required. Contact Quality Monitor (CQM) or monitoring system of the electrode-patient connection, with either alarm or current shut-off for patient safety. Memory for at least 10 programs with their waveforms and power levels. Protection against defibrillator discharges. Convection refrigeration without ventilator.
21 | Displayed parameters | At least: different modes, amperage output, bipolar/monopolar selection, blended option, error status.
22 | User adjustable settings | At least: power level, bipolar/monopolar and able to able to select cot, blend, coag, or more modes
PHYSICAL/CHEMICAL CHARACTERISTICS
23 | Components(if relevant) | Compact design (a standard dimension could be H x W x D: 25 x 30 x 15 cm; variations are acceptable depending on brands and models)
24 | Mobility, portability(if relevant) | Bench top design
25 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
26 | Electrical, water and/or gas supply (if relevant) | Voltage 100 - 240 VAC 50/60 Hz, plug based on country requirement. Protections against over-voltage and over-current line conditions. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage (when needed and/or applicable).
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
27 | Accessories (if relevant) | At least the following accessories ("n.x" before the requirement states the minimum quantity necessary to have a "ready to start" complete equipment): - n.1 smoke evacuator provided with ALL accessories and spare parts necessary to be used with the ESU for example (but not limited to): hoses, pencil adapters, HEPA or ULPA filters, etc. With regard to this device the complete list of accessories and spare parts offered shall be detailed as well as the electrical characteristics (voltage to be compliant with the "utility requirements" previously specified), the type of filters provided, the filters life time in hours and ALL the other technical characteristics. - n.1 equipment Cart (to accomodate the main unit and the smoke evacuator) - n.2 reusable return electrode, adult size, flexible, with connector cable minimum 2 m long - n.2 reusable return electrode, pediatric size, flexible, with connector cable - n.1 reusable handle for monopolar electrodes - n.1 reusable bipolar forceps, sterilizable, straight, short - n.1 reusable bipolar forceps, sterilizable, straight, long - n.1 reusable bipolar forceps, sterilizable, bayonet - n.1 set of reusable electrodes: wired, various sizes and shapes, at a minimum has: 3 -5 mm ball, square loop (smaller), semicircular loop (larger)
28 | Sterilization process for accessories (if relevant) | Not required
29 | Consumables / reagents (if relevant) | At least the following consumables ("n.x" before the requirement states the minimum quantity necessary to have a "ready to start" complete equipment): - n.50 disposable return electrode (two areas) - n.100 pencils with cable - removable electrodes: ball, knife, needle (n.1 package of at least n.40 electrodes per each type). - n.4 Vessel sealing clamp (when necessary)
30 | Spare parts (if relevant) | At least 2-year spare parts kit as per preventive maintenance programme.
31 | Other components (if relevant) | N/A
PACKAGING
32 | Sterility status on delivery (if relevant) | N/A
33 | Shelf life (if relevant) | At least 10 years
34 | Transportation and storage (if relevant) | Unit shall be supplied protectively packed for safe transportation and delivery.
35 | Labelling (if relevant) | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity). Should adhere to the most current version of ISO 15223-1: Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements
ENVIRONMENTAL REQUIREMENTS
36 | Context-dependent requirements | Capable of operating continuously in ambient temperature of 10 to 35 (preferably 40) deg C and relative humidity of 15 to 85 (preferably 90)%, non-condensing. Capable of being stored in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%
TRAINING, INSTALLATION AND UTILISATION
37 | Pre-installation requirements(if relevant) | Not required
38 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation Supplier to perform installation, safety and operation checks before handover
39 | Training of user/s (if relevant) | Training of users in operation and technicians in basic maintenance shall be provided
40 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes. Unit layout to enable easy cleaning and sterilization of all surfaces.
WARRANTY AND MAINTENANCE
41 | Warranty | Not less than 2 years. Specific inclusions and exclusions to be listed.
42 | Maintenance tasks | List of equipment and procedures required for local routine maintenance should be provided.
43 | Type of service contract | Costs and types of post-warranty service contract available should be described (when needed). Contact details of manufacturer, supplier and local service agent to be provided.
44 | Spare parts availability post-warranty | Availability of accesories, consumables and spare parts for at least 7 years. Guaranteed time period of availability of spare parts post-warranty should be pointed out.
45 | Software / Hardware upgrade availability | Provided, when necessary
DOCUMENTATION
46 | Documentation requirements | User / technical manual (including procedures for decontamination) to be supplied in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection to be provided (when applicable). List to be provided of equipment and procedures required for local calibration, if necessary, and routine maintenance. List to be provided of common spares and accessories, with part numbers.
DECOMMISSIONING
47 | Estimated Life Span | Not less than 10 years
SAFETY AND STANDARDS
48 | Standards, for the manufacturer and the equipment | NMPA certificate provided (Chinese Suppliers only). Certified Quality Management System for medical devices (e.g. ISO 13485, or Good Manufacturing Practice - GMP). Application of risk management to medical devices (e.g. ISO 14971). General quality management (e.g. ISO 9001: Quality management systems – Requirements).
49 | Regulatory Approval / Certification | Free Sales Certificate (FSC) provided by any of the following countries: Australia, Canada, Japan, USA and European Community (e.g. FDA and/or CE certificate given by a third certified party for the specific medical devices proposed (no only declaration of conformity).
50 | International standards | Compliance to the following international standards, when applicable, or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended but compliance to previous standards versions could be considered. IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-1:2015 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems. IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. EN IEC 60601-2-2:2018 - Medical electrical equipment - Part 2-2: Particular requirements for the basic safety and essential performance of high frequency surgical equipment and high frequency surgical accessories.