## Foglio1
Blood pressure measurement device
i | Version No. | 1
ii | Date of initial version | 1 December 2019
iii | Date of last modification | 1 December 2019
iv | Date of publication | 31 December 2019
v | Completed / submitted by | WHO Working Group
Name, category and coding
1 | WHO Category / Code | To be determined
2 | Generic name | Electronic blood pressure monitor
3 | Specific type or variation (optional) | Electronic (automated, semi-automated) sphygmomanometer
4 | GMDN name | Automatic-inflation electronic sphygmomanometer, non-portable
5 | GMDN code | 16173
6 | GMDN category | Automatic, electronic, oscillometric
7 | UMDNS name | Sphygmomanometers, electronic, automatic. Sphygmomanometers, electronic, automatic, oscillometric monitors
8 | UMDNS code | 18326, 25209
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Non-invasive BP monitors; oscillometric sphygmomanometers; oscillotonometers; spot check monitors; spot checking; sphygmomanometer, automatic
11 | Alternative code/s (optional)
12 | Keywords (optional) | Automatic electronic sphygmomanometers non-invasive. Digital automatic non-invasive BP monitor
13 | GMDN/UMDNS definition (optional) | An electrically powered device designed to non-invasively measure BP, with a self-contained software program to regulate automatic arm-cuff inflation and measurement cycles. It typically displays current heart rate and mean arterial pressure in addition to systolic and diastolic BP; it may have memory to store values and may sound an alarm if BP exceeds pre-set limits. This device is not designed to be portable and is typically used at the bedside.
Purpose of use
14 | Clinical or other purpose | Physical examination; diagnosis of hypertension; monitor, measure and display arterial blood pressure
15 | Level of use (if relevant) | Ambulatory care centre, health centre, district hospital, provincial hospital, specialized hospital, home
16 | Clinical department/ ward(if relevant) | All areas
17 | Overview of functional requirements | The main unit includes controls and displays numerical data for BP. It also includes appropriate attached cuffs (probes, and sensors, depending on their configuration) that allow sequential, periodic and/or simultaneous measurements.
Technical characteristics
18 | Detailed requirements | Measurement ranges: systolic (mm Hg), 60–250, 290 preferred for adults, 30–160 for children and 20–120 for neonates. Diastolic (mm Hg), 30–180 adults, 10–150 paediatric, 10100 neonate. Mean arterial pressure (mm Hg), 30–250 adults, 30–160 children, 30–110 neonates. Pulse (beats per min), 30–150 adult and children, 30–180 neonates. Inflation pressure (mm Hg) 150–260 adults, 85–140 neonates; adjustable or automatically set preferred. Auto deflate pressure (mm Hg), 300 adults, 150 neonates. Measurement interval, min: User selectable: ≥ 5 choices. Cuff sizes: neonatal, paediatric, adult, large adult, thigh. Measurement time (s) ≤ 60, user selectable. Automatic 0 required. Display may include tabular and/or graphic trends (user preference). Equipment alarms required: cuff leak, cuff disconnect, failure to take successful reading, low-battery notice. Equipment alarms preferred: hose leak, inflation or deflation error. Sphygmomanometer should automatically deflate if the cuff pressure reaches 300 mm Hg for an adult and 150 mm Hg for a neonate.
19 | Displayed parameters | The unit should display the following numerical values: systolic pressure, diastolic pressure, pulse rate and mean arterial pressure. Other parameters are optional. The unit should alert the operator, either visually or audibly.
20 | User adjustable settings | Inflation pressure should be adjustable or automatically set according to a previous or current pressure reading or individual requirements. Time between automatic BP measurement cycles should be selectable from at least five values over a range of 1 to 60 min. Set alarm volume and limits within the specified measurement ranges.
Physical/chemical characteristics
21 | Components(if relevant) | Rubber tubes to be detachable from other parts, allowing periodic cutting of decayed ends. Gauge body to include clip for mounting on cuff. Tube length to be > 30 cm. Different cuff sizes available (small or neonate, medium or paediatric, large or adult and extra-large or large adult). Cuff material to be removable and washable.
22 | Mobility, portability(if relevant) | Wall, portable, table-top, mobile stand
23 | Raw Materials(if relevant) | N.A.
Utility requirements
24 | Electrical, water and/or gas supply (if relevant) | AC: 120/240, 50/60 Hz DC: Rechargeable battery (for at least 1 h of operation, single-use or rechargeable)
Accessories, consumables, spare parts, other components
25 | Accessories (if relevant) | Mobile stand
26 | Sterilization process for accessories (if relevant) | N.A.
27 | Consumables / reagents (if relevant) | Single-use cuffs in the following sizes: neonatal (10–15 cm), paediatric (14–22 cm), adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). The sizes of the cuffs depend on the manufacturer but should not deviate by ± 5 cm from the stated sizes. Batteries
28 | Spare parts (if relevant) | Rubber tube (length > 30 cm), reusable cuffs in the following sizes: neonatal (10–15 cm), paediatric (14–22 cm), adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). The sizes of the cuffs depend on the manufacturer but should not deviate by ± 5 cm from the stated sizes. Tubing, valve
29 | Other components (if relevant) | Protective Case
Packaging
30 | Sterility status on delivery (if relevant) | Single-use cuffs must be delivered sterile.
31 | Shelf life (if relevant) | Minimum shelf life for single-use cuffs must be 1 year from the date of reception.
32 | Transportation and storage (if relevant) | Storage environment humidity: 10–95% relative humidity. Storage environment temperature: –20 to 60 °C
33 | Labelling (if relevant) | With the proper certification and validation requested, plus those required in each country
Environmental requirements
34 | Context-dependent requirements | Handling environment temperature: –20 to 60 °C
Training, installation and utilization
35 | Pre-installation requirements(if relevant) | N.A.
36 | Requirements for commissioning (if relevant) | Battery, uninterruptable power source, appropriate cuffs
37 | Training of user/s (if relevant) | All users (physicians nurses, other medical staff) shall have initial training in operation. Biomedical or clinical engineer or technician, medical staff, manufacturer or servicer shall have initial training in operation and basic maintenance by manufacturer, and subsequently if necessary.
38 | User care(if relevant) | Clean surface of device and wash reusable cuffs as stated by manufacturer.
Warranty and maintenance
39 | Warranty | 2 years
40 | Maintenance tasks | Cables and lead wires should be inspected periodically for breaks and cracks.
41 | Type of service contract | N.A.
42 | Spare parts availability post-warranty | 5 years after discontinuation by factory
43 | Software / Hardware upgrade availability | Software upgrade required and if available from factory
Documentation
44 | Documentation requirements | User, troubleshooting and service manuals must be available to the client, preferably in the national language(s) and/or in another language authorized by the national regulatory agency. Certificate of calibration and validation to be provided. List of equipment and procedures required for local calibration and routine maintenance to be provided List of important spares and accessories, with their part numbers and cost, to be provided. Contact details of manufacturer, supplier and local service agent to be provided.
Decommissioning
45 | Estimated Life Span | 10 years.
Safety and standards
46 | Standards, for the manufacturer and the equipment | Depends on the country. Examples: Class A (Global Harmonization Task Force Rule 4); Class II (USA); Class I (Australia, Canada and Japan); Class IIa (European Union)
47 | Regulatory Approval / Certification | Proof of regulatory compliance (e.g. registration, clearance, approval) must be provided as appropriate per the product’s risk classification, by regulatory agency (e.g. by a founding member of IMDRF-EU, USA, Canada, Australia, Japan). Else approved by local national regulatory agency.
48 | International standards (for reference) | Standards applicable to the product and to the manufacturing process are listed below. Compliance to the last available version of the international standard or to its local equivalent standard is recommended and proof of compliance must be provided. Non-exhaustive list of standards applicable to general quality systems for medical devices and specific for BPMD: • ISO 13485:2016, Medical devices – Quality management systems – Requirements for regulatory purposes • EN ISO 14971:2012, Medical devices – Application of risk management to medical devices • ISO 14155:2011, Clinical investigation of medical devices for human subjects – Good clinical practice • ISO 14971:2007, Medical devices – Application of risk management to medical devices • IEC 80601-2-30:2018 Medical electrical equipment – Part 2-30: Particular requirements for basic safety and essential performance of automated non-invasive sphygmomanometers • ISO 16142-1:2016, Medical devices – Recognized essential principles of safety and performance of medical devices – Part 1: General essential principles and additional specific essential principles for all non-IVD medical devices and guidance on the selection of standards Non-exhaustive list of standards applicable to electronic BP devices: • AAMI/ESH/ISO 81060 Universal Standard for the Validation of Blood Pressure Measuring Devices Noninvasive phygmomanometers – Part 2: Clinical investigation of automated measurement type • ISO 81060-2:2018(E) Non-invasive sphygmomanometer standard Part 2: Clinical investigation of intermittent automated measurement type • ISO/IEEE 11073-10407:2010 (Part 10407: Device specialization – Blood pressure monitor) • IEC 80601-2-30:2009 (Part 2-30: Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometers) • DS/EN 1060-3 Non-invasive sphygmomanometers – Part 3: Electro-mechanical blood pressure measuring system
49 | Regional/local standards | ANSI/AAMI SP10:2002 & ANSI/AAMI SP10:2002/A1:2003 (Manual, electronic or automated sphygmomanometers) DS/EN 1060-3 Non-invasive sphygmomanometers - Part 3: Electro-mechanical blood pressure measuring system GOST R 50267.30 Medical electrical equipment. Part 2. Particular requirements for safety of automatic cycling indirect blood pressure monitoring equipment JIS T 1115:2005 Non-invasive automated sphygmomanometers
50 | Regulations | Compliance with (where applicable, but not limited to, and last available version): USA: CFR - Code of Federal Regulations, Title 21, Part 820 CFR - Code of Federal Regulations, Title 21, Part 870, Section 1130 / Non-invasive blood pressure measurement system Japan: MHLW Ordinance No. 16916156000 Aneroid sphygmomanometer European Commission: Council Directive 93/42/EEC of 14 June 1993 on Medical Devices Regulation (EU) 2017/745 of the European Parliament and the Council on Medical Devices