## Foglio1
Oxygen concentrator, flowsplitter for newborn/child (The specification detailed is for general reference only and it refers to the "Flow splitter, 5 flowmeters 0-2 L/min, for paediatric use")
Published on: 19 November 2020 - Version: 2
Version 1: September 2019 (WHO-UNICEF)
Technical Specification | Minimum requirements
General technical requirements | Flow splitter from a single or double oxygen supply. Equipped with four or five independent, pressure-compensated, Thorpe tube flowmeters, to regulate the flow of medical gas. Suitable for tabletop and/or wall mount. Flowmeters capacity: 0.125–2 L/min. Accuracy: ± 10%. Inlet port to be compatible with all the international standards for oxygen fittings, including DISS, threaded and non-threaded, 6 mm barbed – availability of different ports and/or adapters to be stated. 6 mm barbed outlet as standard – availability of adapters and outlet options to match all the international standards for oxygen fittings to be stated. Transparent, clearly readable and graduated (metric system) column, shatter-resistant polymer certified for medical use. Needle valve and body constructed of brass or aluminium. Adjustment knobs to have rough surface to prevent slipping. Colour-coded flowmeter preferable, e.g. to ISO 32. Flowmeter stand hard plastic or metal epoxy painted. Capable to be disinfected with hospital grade detergents.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity). Gas type, calibration temperature and pressure should be specified on the label.
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485, or good manufacturing practice [GMP]). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: Colour coding ISO or ANSI for medical gases. Conforms to ISO, NFPA, and/or CGA standards, and/or UL or CSA approved. ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen. ISO 15002 Flow-metering devices for connection to terminal units of medical gas pipeline systems. ISO 18562 Biocompatibility evaluation of breathing gas pathways in healthcare applications. ISO 10524 Pressure regulators for use with medical gases. ISO 18082 Anaesthetic and respiratory equipment – Dimensions of non-interchangeable screw-threaded (NIST) low-pressure connectors for medical gases. ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements. ISO 5359 Low-pressure hose assemblies for use with medical gases. ISO 32 Colour coding for medical gases.
Warranty | 2 years with regards efficiency and quality of the product. Availability of accessories and spare parts for at least 2 years.