## Foglio1
Video-laryngoscope
i | Version No. | 1
ii | Date of initial version | 2020-06-23
iii | Date of last modification | 2020-07-22
iv | Date of publication | 2020-11-19
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | Under development
2 | Generic name | Video-Laryngoscope, rigid
3 | Specific type or variation (optional) | Bladed video laryngoscope
4 | CND code (https://ec.europa.eu/health/md_topics-interest/overview_en) | Z12021004
5 | CND nomenclature | VIDEOLARYNGOSCOPES
Purpose of use
6 | Clinical or other purpose | To manipulate the tongue and enable a clear view of the trachea for surgical/mechanical ventilation/ intubation procedures.
7 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
8 | Clinical department/ ward(if relevant) | Outpatients, ear, nose and throat (ENT), operating theatre, emergency room (ER), intensive care unit (ICU)
9 | Overview of functional requirements | Video laryngoscope rigid equipment used to establish a direct line-of-sight between the practitioner’s eye and the larynx using a digital monitors. The system is composed by an handle, reusable or single use blades with included a video-camera function and a monitor/display mounted on the device’s handle.
Technical characteristics
10 | Detailed requirements | Safe intubation of paediatrics and adults. Highly curved d-blade and paediatric d-blade or better for particularly difficult airway. TFT/LCD or equivalent or better technology. Colour display/monitor attached to the handle or built-in or external. If the display is attached to the handle, the dimension should be not less than 2 inch; if the monitor is external the size should be not less than 7 inch. High definition video processor. With white balance adjustment, preferable. Quick-shot camera button which makes it easy to operate or similar system for digital still images recording. Sizes: neonates (preferable), paediatric and adult. CMOS/CCD full-colour camera with integrated LED or better light source.
11 | Displayed parameters | N/A
12 | User adjustable settings | User shall be able to vary intensity of light in addition to powering on/off.
Physical/chemical characteristics
13 | Components(if relevant) | The system/equipment provided is composed at least by the following components/parts: handheld unit; on/off switch to be robust and easy to use; monitor/display integrated or external; compatible set of reusable or disposable blades; protective case for clean storage and safe transport. Unit case should be hard and splashproof. There must be no sharp edges on the unit.
14 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand.
15 | Raw Materials(if relevant) | Reusable blades or disposable blades to be made of surgical grade material or better quality material.
Utility requirements
16 | Electrical, water and/or gas supply (if relevant) | Powered by internal, rechargeable, replaceable battery. Battery cover to be secure but simple to open and battery compartment to be sealed against liquid ingress. Battery charger to operate from input supply 110–220 V, 60–50 Hz, ± 10% (battery charger built-in, preferably, or external) with protection against over-voltage and over-current line conditions.
Accessories, consumables, spare parts, other components
17 | Accessories (if relevant) | The equipment is provided complete of the set of accessories and/or consumables for adult and paediatric, depending on which one of the following alternative configurations requested will be offered: Configuration 1) At least three compatible Macintosh blades of different measures (adult and paediatric) and at least two compatible Miller reusable blades (adult and paediatric). OR Configuration 2) One reusable blade with an adjustable length to be used with different adult and paediatric patients; in this case the blade should be provided with a complete set of disposable blade covers or tips. OR Configuration 3) Complete set of disposable blades with at least three different measures for adult and paediatric, and compatible with the system provided. Protective case for clean storage and safe transport (with dedicated space to transport at least: the handle, the monitor, a set of batteries and 2 blades, when applicable).
18 | Sterilization process for accessories (if relevant) | Supplier to describe any sterilization process required for accessories, if any.
19 | Consumables / reagents (if relevant) | Tongue depressors, single use. Complete set for adult and paediatric of fully compatible single-use blades (or blade covers) (if necessary to operate and when applicable). Supplier to describe any additional necessary consumables, detailing shelf life and number of uses.
20 | Spare parts (if relevant) | Replacement battery pack, supplied empty of charge. Spare light bulb compatible with the device provided.
21 | Other components (if relevant) | N/A
Packaging
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | N/A
24 | Transportation and storage (if relevant) | Unit shall be supplied protectively packed for safe onward shipping. Labelling on the primary packaging under local regulations.
25 | Labelling (if relevant) | N/A
Environmental requirements
26 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0–50 °C and relative humidity of 15–85%, preferably 90%. Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85%, preferably 90%.
Training, installation and utilization
27 | Pre-installation requirements(if relevant) | Not required.
28 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation. Supplier to perform installation, safety and operation checks before handover.
29 | Training of user/s (if relevant) | Training of users in operation and technicians in basic maintenance shall be provided.
30 | User care(if relevant) | The whole unit is to be cleanable with standard disinfectants usually used in hospitals.
Warranty and maintenance
31 | Warranty | Not less than 2 years with specific inclusions and exclusions to be clearly listed. Contact details of manufacturer, supplier and local service agent to be provided.
32 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance.
33 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
34 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
35 | Software / Hardware upgrade availability | Software updates included during warranty. Upgradable device with additional programs and software fuctions (when available). Guaranteed time period of support and updates (when necessary) availability post-warranty shall be described.
Documentation
36 | Documentation requirements | User/technical manual to be supplied in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection to be provided, when applicable. List to be provided of equipment and procedures required for local calibration, if necessary, and routine maintenance. Battery disposal according local laws.
Decommissioning
37 | Estimated Life Span | Not less than 5 years.
Safety and standards
38 | Standards, for the manufacturer and the equipment | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (e.g. ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 Medical devices – Application of risk management to medical devices).
39 | Regulatory Approval / Certification | Free sales certificate (FSC). Certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
40 | International standards | Compliance to the following international standards, when applicable, or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended but compliance to previous standards versions could be considered. IEC 60601-1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2007 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 7376-1:1994 Laryngoscopic fittings – Part 1: Conventional hook-on type handle-blade fittings IEC 60601-2-18 Ed. 3.0:2009 (b) Medical electrical equipment – Part 2-18: Particular requirements for the basic safety and essential performance of endoscopic equipment. ISO 8600-1:2013 Endoscopes – Medical endoscopes and endotherapy devices – Part 1: General requirements. ISO 8600-4:1997 Optics and optical instruments – Medical endoscopes and certain accessories – Part 4: Determination of maximum width of insertion portion ISO 8600-7:2012 Endoscopes – Medical endoscopes and endotherapy devices – Part 7: Basic requirements for medical endoscopes of water-resistant type.