## Foglio1
Blood gas analyser, automated/semi-automated (portable device, including cartridges and solutions)
i | Version No. | 2
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2020-07-22
iv | Date of publication | 2020-11-19
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | under development
2 | Generic name | Arterial whole blood pH and gases (pO2, pCO2) analyser
3 | Specific type or variation (optional) | Blood
4 | CND code (https:// ec.europa.eu/health/ md_topics-interest/ overview_en) | W0201069002
5 | CND nomenclature | BLOOD GAS PORTABLE ANALYSERS
Purpose of use
6 | Clinical or other purpose | Measurement of arterial blood gases and pH. Blood gas measurements are requested in respiratory, metabolic or kidney disease, especially if there is respiratory distress, to evaluate the amount of oxygen and the acid/base balance.
7 | Level of use (if relevant) | Health post, health centre, district hospital, provincial hospital, specialized hospital and outreach (mobile clinics)
8 | Clinical department/ ward(if relevant) | Emergency room (ER), neonatal internsive care unit (NICU), surgery, outpatient, intensive care unit (ICU), hospital
9 | Overview of functional requirements | The equipment tests arterial whole blood samples or capillary whole blood samples; must be anti- coagulated (balanced lithium heparin). Blood gas/pH analysers use reagents kits or specific consumables or are cartridge-based systems. They have sensors for assessing pH, pO2, and pCO2, in anticoagulated blood samples. Once a proprietary cartridge is inserted into a point-of-care (POC) blood gas/pH device or a specific reagent kit is used, the pH, pO2, and pCO2 values are reported. Temperature changes affect measured values, test cartridges in particular are inserted into a temperature-controlled block within a POC blood gas/pH analyser that is maintained at 37 °C (normal body temperature). Cartridges are disposed as biohazardous garbage after testing.
Technical characteristics
10 | Detailed requirements | Point-of-care equipment, easy to be moved between different hospital departments. Analytical performance/accuracy should remain stable throughout the measurement of the analyte and can be monitored using internal quality controls (ICQ) which are performed automatically by the analyser; frequency will depend on manufacturer. The assay is linear across at least the measuring ranges: pH 6.75–7.85; pO2 1.33–73.3 kPa; pCO2 1.6–14.7 kPa. Analytical specificity (interferences): Method should be evaluated and demonstrate no interference from ascorbate, bilirubin, haemoglobin, lipid, acetaminophen, acetylsalicylic acid. Blood gas analysis should have bias from the reference method not higher than 5%. Potentiometric measurement principle on a dedicated blood gas analyser, quality equivalent measurement principles can be considered. The equipment should have a LCD (or equivalent technology) monitor/display to easy visualize data and to manage parameters settings. Sample volume will vary by instrument but is typically not higher than 195 µl for a complete assay. Values reported are quantitative, results should be reported at least in pH units, kilo Pascals (kPa; O2 and CO2) and mmol/L (bicarbonate). Bicarbonate is calculated from the Henderson-Hasselbalch equation. Availability of other additional and selectable pressure measurement units, is preferable. Time to result not higher than 3 minutes. Specimen throughput not lower than 20 specimens per 8 hour day. The manufacturer should precisely state the tests limitations, if any. Quality control checks will be automatically performed at intervals defined by manufacturer. External quality assessment (EQA) material should be provided with the equipment (when necessary, and available). The equipment should use specific/proprietary or preferably non-branded (commonly available in the market) test kits/reagents or dedicated cartridges. The shelf life of kit/reagents or cartridges upon delivery should not be less than 6 months. The equipment should be provided with automatic calibration function. The equipment should be provided with a built-in or external compatible printer. Equipment preferably provided with bar-code reader for both reagents/kits/cartridges and samples. The equipment should have the capability of data storage (at least 300 patient results and 50 control results) and transfer to external PC through an interface (hardware and software necessary for data transfer should be provided with the equipment). Equipment provided with all components, consumables and accessories necessary for having a “ready to start” device and for performing the required quality control (QC) and calibration tests.
11 | Displayed parameters | At least the following paramentes displayed: PO2, PCO2 and pH.
12 | User adjustable settings | Depending on the manufacturer and models.
Physical/chemical characteristics
13 | Components(if relevant) | Equipment supplied with dedicated case for safe transport (preferably, when available). Equipment supplied with a dedicated dust cover (preferably, when available).
14 | Mobility, portability(if relevant) | There must be no sharp edges on the unit.
15 | Raw Materials(if relevant) | Easy and safe to be transported from different departments.
Utility requirements
16 | Electrical, water and/or gas supply (if relevant) | Powered by mono-phase electrical source, AC line 110–220 V, 60–50 Hz, ± 10% with line connection plug compatible with local standards. Power should preferably be delivered via an online uninterruptable power supply (UPS) providing the same output as local electric supply. Protections against over-voltage and over-current line conditions. Powered preferably also by internal, rechargeable, replaceable battery. Battery capacity at least 1 hour of battery power when analyser is in “ready mode”, or approximately 25 samples. Battery charger to operate from input supply 110–220 V, 60–50 Hz, ± 10% (battery charger built-in or external).
Accessories, consumables, spare parts, other components
17 | Accessories (if relevant) | Trolley or cart with brakes, able to carry the analyser and his UPS (when necessary, preferable). Equipment provided with any accessory or dedicated device necessary to the proper functioning and utilization of the equipment included (e.g. batteries, mounting kit, first test kit).
18 | Sterilization process for accessories (if relevant) | N/A
19 | Consumables / reagents (if relevant) | Complete reagent kits /cartridges with replacement electrodes if bundled. Complete self-calibration and quality controls kits/cartridges.
20 | Spare parts (if relevant) | All reagents and consumables should be provided by manufacturer; components should be clearly listed.
21 | Other components (if relevant) | Replacement battery pack, supplied empty of charge.
Packaging
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | N/A
24 | Transportation and storage (if relevant) | Unit shall be supplied protectively packed for safe onward shipping. Labelling on the primary packaging under local regulations.
25 | Labelling (if relevant) | N/A
Environmental requirements
26 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0–50 °C and relative humidity of 15–85%, preferably 90%. Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85%, preferably 90%.
Training, installation and utilization
27 | Pre-installation requirements(if relevant) | Not required
28 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation. Supplier to perform installation, safety and operation checks before handover.
29 | Training of user/s (if relevant) | Training of users in operation and technicians in basic maintenance shall be provided.
30 | User care(if relevant) | The whole unit is to be cleanable with standard disinfectants usually used in hospitals.
Warranty and maintenance
31 | Warranty | Not less than 2 years with specific inclusions and exclusions to be clearly listed. Contact details of manufacturer, supplier and local service agent to be provided.
32 | Maintenance tasks | List of equipment and procedures required for local calibration and routine maintenance should be provided.
33 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
34 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
35 | Software / Hardware upgrade availability | Software updates Included during warranty. Guaranteed time period of support and updates (when necessary) availability post-warranty shall be described.
Documentation
36 | Documentation requirements | User/technical manual to be supplied in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection to be provided, when applicable. List to be provided of equipment and procedures required for local calibration, if necessary, and routine maintenance. Battery disposal according local laws.
Decommissioning
37 | Estimated Life Span | Not less than 5 years
Safety and standards
38 | Standards, for the manufacturer and the equipment | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (e.g. ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 Medical devices – Application of risk management to medical devices).
39 | Regulatory Approval / Certification | Free sales certificate (FSC). Certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
40 | International standards | Compliance to the following international standards, when applicable, or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended. ISO 23640: In vitro diagnostic medical devices – Evaluation of stability of in vitro diagnostic reagents. ISO 18113-1: In vitro diagnostic medical devices – Information supplied by the manufacturer (labelling) – Part 1: Terms, definitions and general requirements.