## Foglio1
Computerized treatment planning systems (three- dimensional): external beam treatment planning system (TPS)
i | Version No. | 1
ii | Date of initial version | 10 October 2019
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | Working group (WHO/IAEA)
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Treatment planning system
3 | Specific type or variation (optional) | Not applicable
4 | GMDN name | Radiation therapy treatment planning computer; Radiation therapy treatment planning workstation
5 | GMDN code | 40887; 40860; 40941
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Computers, Radiotherapy Planning System
8 | UMDNS code | 13281
9 | UNSPS code (optional) | Not applicable
10 | Alternative name/s (optional) | Not applicable
11 | Alternative code/s (optional) | Not applicable
12 | Keywords (optional) | External beam radiotherapy
13 | GMDN/UMDNS definition (optional) | (40887) An application software programme intended to facilitate/support the delivery of therapeutic radiation by enabling computer-assisted functionality in one or more of the following areas: treatment planning, simulation, record/verify, beam block computer-aided design/computer-aided manufacturing (CAD/CAM), quality assurance, and patient positioning. Some types may be combined with specific hardware or firmware accessories or configurations in order to function as intended. (40860) A dedicated mainframe computer, personal computer (PC) or PC-based platform and associated hardware, firmware, and operating system software used specifically to control and monitor the operation of a radiation therapy treatment planning system (RTTPS) and associated processing, display and analysis functions. (40941) A freestanding image processing device that is a component of radiation therapy treatment planning configuration. It is sometimes considered to be a component of a picture archiving and communication system (PACS) or a local area network (LAN) configuration, e.g. a radiology information system (RIS) or hospital information system (HIS). It does not contain any controls for the direct operation of a radiation therapy system. It can receive and transmit data both online and offline and is configured to provide the capability to remotely process, manipulate and view patient images and information used in treatment planning, record and verification.
14 | EMDN code | Z11010482
15 | EMDN nomenclature | RADIOTHERAPY PLAN EVALUATION INSTRUMENTS - SOFTWARE
Purpose of use
16 | Clinical or other purpose | Preparation of patient treatment plans for external beam radiotherapy (EBRT).
17 | Level of use (if relevant) | Hospital
18 | Clinical department/ ward(if relevant) | Radiation Oncology Department
19 | Overview of functional requirements | Software to generate treatment plans for EBRT. The software shall include modules for image import and registration, contouring of target volumes and organ at risk, planning of treatment beams, dose calculation, plan review and export, and beam modelling.
Technical characteristics
20 | Detailed requirements | • Treatment planning workstations, including dual 23-inch monitors, keyboard, mouse and network capability. • A module to allow import of patient data sets from various imaging modalities that are used to facilitate target definition, using the Digital Imaging and Communications in Medicine (DICOM) standard. Image import may be achieved through direct connectivity to the radiology PACS system. • An image registration module using either rigid and/or deformable registration modes. • 3D visualization capability for patient data display, beam display and dose distribution display. • Contouring tools that allow the definition in 3D of structures, including target, organs at risk and patient outline. • Automated tools to allow the expansion of the clinical target volume (CTV) to a planning target volume (PTV) with non-uniform margins in three dimensions. • The ability to add bolus structures to the patient data set of various shape and density. • A comprehensive “forward planning” environment which allows the user to modify beam weights, beam positioning, jaw position, wedges and blocks or multileaf collimator (MLC) to optimize the treatment plan. • Photon beam and electron beam algorithms which calculate the dose to the patient considering the 3D nature and heterogeneity of the patient data set. The photon beam algorithm shall use advanced methods (e.g. convolution, Boltzmann transport, Monte Carlo). The electron beam algorithm shall be based on pencil beam or Monte Carlo methods. The dose matrix (calculation grid) should be user adjustable and the accuracy of the dose calculations shall be such that it is possible to meet the tolerance criteria in Table 4 of IAEA TECDOC 1540. • The ability to calculate MU for all beam models. • The ability to allow dose prescription to a point, volume or isodose line. • Advanced plan review and evaluation tools, including dose volume histograms (DVHs), dose statistics, 2D and 3D dose visualization, and plan addition and plan comparison. • The ability to generate and view digitally reconstructed radiographs (DRRs). • Display of all relevant treatment unit parameters in the IEC 61217 scale convention as at least one option. • A comprehensive beam modelling module that allows the configuration of complete geometric and dosimetric models for treatment unit photon and electron beams. The module shall have the following features: o ability to import measured beam profiles and output factors; o ability to model dynamic, fixed and internal wedges; o tools to allow the comparison of the beam model and measured data; and o security features which protect beam data and beam models from modification. • A module to allow the creation of CT number to mass density or electron density data for various CT scanners for use by the photon and electron beam algorithms. • User and password security that allows approval/locking of treatment plans and different levels of access to the functionality of the TPS based on the user’s profile e.g. administrator, planner, medical physicist, physician. • A printer for A3/4 output of isodose distributions, beam shapes and treatment plan parameters. • A module to allow export of beam block shapes to a third party block cutting device. • A module to allow export of approved treatment plans and DRRs to an oncology information system (OIS).
21 | Displayed parameters | Not applicable
22 | User adjustable settings | Not applicable
Physical/chemical characteristics
23 | Components(if relevant) | Major components of the TPS are networked computer workstations, server (if applicable), printer and software.
24 | Mobility, portability(if relevant) | Fixed, not mobile
25 | Raw Materials(if relevant) | Not applicable
Utility requirements
26 | Electrical, water and/or gas supply (if relevant) | Single-phase electrical power and air-conditioning.
Accessories, consumables, spare parts, other components
27 | Accessories (if relevant) | Not applicable
28 | Sterilization process for accessories (if relevant) | Not applicable
29 | Consumables / reagents (if relevant) | Not applicable
30 | Spare parts (if relevant) | Not applicable
31 | Other components (if relevant) | Not applicable
Packaging
32 | Sterility status on delivery (if relevant) | Not applicable
33 | Shelf life (if relevant) | Not applicable
34 | Transportation and storage (if relevant) | If storage is required prior to installation, the storage area should be secure, clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof.
35 | Labelling (if relevant) | Not applicable
Environmental requirements
36 | Context-dependent requirements | It is the responsibility of the hospital to ensure that the manufacturer’s recommended operation and storage conditions, as given in their site planning guide, are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
Training, installation and utilization
37 | Pre-installation requirements(if relevant) | Network availability
38 | Requirements for commissioning (if relevant) | Acceptance testing, commissioning of radiation beam models and dose calculation algorithms, baselines for quality control.
39 | Training of user/s (if relevant) | Applications training in configuration and clinical operation.
40 | User care(if relevant) | The equipment shall only be used in accordance with departmental policies, procedures and work instructions, and consistent with manufacturer’s use instructions, to ensure safety and protection.
Warranty and maintenance
41 | Warranty | At least 12 months
42 | Maintenance tasks | Software updates as required
43 | Type of service contract | Service contract, including software upgrades and access to helpdesk.
44 | Spare parts availability post-warranty | 4 years minimum
45 | Software / Hardware upgrade availability | To be specified by manufacturer.
Documentation
46 | Documentation requirements | Comprehensive user manual, including emergency procedures.
Decommissioning
47 | Estimated Life Span | 5 years
Safety and standards
48 | Risk classification | US FDA: Device Class 2
49 | Regulatory approval/certification | Radiation regulatory approval for equipment possession, use and premises.
50 | International standards | IEC, Medical Electrical Equipment, Part 1: General Requirements for Safety, Rep. IEC 60601-1:2005+AMD1:2012. IAEA, Radiation Protection and Safety of Radiation Sources, Basic Safety Standards, No. GSR Part 3, 2014. IEC, Medical electrical equipment – Requirements for the safety of radiotherapy treatment planning systems, IEC 62083:2009. IEC, Radiotherapy equipment – Coordinates, movements and scales, IEC 61217, 2011. IAEA, Specification and acceptance testing of radiotherapy treatment planning systems, TECDOC 1540, 2007.
51 | Regional/local standards | Country-specific and regional standards may apply.
52 | Regulations | USA: 21CFR892.5050 Medical charged-particle radiation therapy system.
Computerized treatment planning systems (three- dimensional): brachytherapy treatment planning system (TPS) technical specifications
i | Version No. | 1
ii | Date of initial version | 10 October 2019
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | Working group (WHO/IAEA)
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Treatment planning system
3 | Specific type or variation (optional) | Not applicable
4 | GMDN name | Remote-afterloading brachytherapy system (component 6)
5 | GMDN code | 38300
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Computers, Radiotherapy Planning System
8 | UMDNS code | 13281
9 | UNSPS code (optional) | Not applicable
10 | Alternative name/s (optional) | Not applicable
11 | Alternative code/s (optional) | Not applicable
12 | Keywords (optional) | Brachytherapy
13 | GMDN/UMDNS definition (optional) | Not applicable
14 | EMDN code | Z11010482
15 | EMDN nomenclature | RADIOTHERAPY PLAN EVALUATION INSTRUMENTS - SOFTWARE
Purpose of use
16 | Clinical or other purpose | Preparation of patient treatment plans for brachytherapy.
17 | Level of use (if relevant) | Hospital
18 | Clinical department/ ward(if relevant) | Radiation Oncology Department
19 | Overview of functional requirements | Software to generate treatment plans for brachytherapy. The software shall include modules for image import and registration, contouring of target volumes and organ at risk, planning of source positions and dwell times, dose calculation, plan review and export, and source modelling. The software shall enable transfer of treatment plans to the brachytherapy treatment unit console.
Technical characteristics
20 | Detailed requirements | • A networked workstation with 23-inch monitor and significant storage capacity. • Software for dose optimization based on points, planes or volumes based on catheter and source dwell positions obtained by reconstruction from isocentric, orthogonal and semi-orthogonal images. • Hardware, software and licenses which allow Digital Imaging and Communications in Medicine (DICOM) compliant image transfer from a computed tomography (CT) scanner, magnetic resonance imaging (MRI) scanner and conventional simulator or a C-arm fluoroscopic X-ray unit. • A flatbed film scanner (or equivalent) to allow the manual input of films or digitized images. • Reference data sets for the source all applicators offered (library) and for the source. • Hierarchical password-controlled security access to source calibration, modelling parameters, library of standard treatment plans and synchronization of source data with the treatment control panel. • System manager software for database maintenance and archiving. • A printer for A3/4 output of isodose distributions and treatment plan parameters. • Software for automatic online transfer of patient and authorized treatment plan data to the treatment afterloader control panel.
21 | Displayed parameters | Not applicable
22 | User adjustable settings | Not applicable
Physical/chemical characteristics
23 | Components(if relevant) | Major components of the TPS are networked computer workstations, printer and software.
24 | Mobility, portability(if relevant) | Fixed, not mobile
25 | Raw Materials(if relevant) | Not applicable
Utility requirements
26 | Electrical, water and/or gas supply (if relevant) | Single-phase electrical power and air-conditioning.
Accessories, consumables, spare parts, other components
27 | Accessories (if relevant) | Not applicable
28 | Sterilization process for accessories (if relevant) | Not applicable
29 | Consumables / reagents (if relevant) | Not applicable
30 | Spare parts (if relevant) | Not applicable
31 | Other components (if relevant) | Not applicable
Packaging
32 | Sterility status on delivery (if relevant) | Not applicable
33 | Shelf life (if relevant) | Not applicable
34 | Transportation and storage (if relevant) | If storage is required prior to installation, the storage area should be secure, clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof.
35 | Labelling (if relevant) | Symbols used according to ISO-15233 standards.
Environmental requirements
36 | Context-dependent requirements | It is the responsibility of the hospital to ensure that the manufacturer’s recommended operation and storage conditions, as given in their site planning guide, are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
Training, installation and utilization
37 | Pre-installation requirements(if relevant) | Network availability
38 | Requirements for commissioning (if relevant) | Acceptance testing, commissioning of source model, baselines for quality control.
39 | Training of user/s (if relevant) | Applications training in configuration and clinical operation.
40 | User care(if relevant) | The equipment shall only be used in accordance with departmental policies, procedures and work instructions, and consistent with manufacturer’s use instructions, to ensure safety and protection.
Warranty and maintenance
41 | Warranty | At least 12 months
42 | Maintenance tasks | Software updates as required
43 | Type of service contract | Service contract, including software upgrades and access to helpdesk.
44 | Spare parts availability post-warranty | 4 years minimum
45 | Software / Hardware upgrade availability | To be specified by manufacturer.
Documentation
46 | Documentation requirements | Comprehensive user manual, including emergency procedures.
Decommissioning
47 | Estimated Life Span | 5 years
Safety and standards
48 | Risk classification | USA FDA: Device Class 2
49 | Regulatory approval/certification | Radiation regulatory approval for equipment possession, use and premises.
50 | International standards | IEC, Medical Electrical Equipment, Part 1: General Requirements for Safety, Rep. IEC 60601-1:2005+AMD1:2012. IAEA, Radiation Protection and Safety of Radiation Sources, Basic Safety Standards, No. GSR Part 3, 2014. IEC, Medical electrical equipment - Requirements for the safety of radiotherapy treatment planning systems, IEC 62083:2009.
51 | Regional/local standards | Country-specific and regional standards may apply.
52 | Regulations | USA: 21CFR892.5700 Remote-controlled radionuclide applicator system.