## Foglio1
Colposcope
i | Version No. | 1
ii | Date of initial version | December 2019
iii | Date of last modification
iv | Date of publication | February 2020
v | Completed / submitted by | WHO
Name, category and coding
1 | WHO Category / Code | Colposcopes, XD2AZ1
2 | Generic name | Colposcope
3 | Specific type or variation (optional)
4 | GMDN name
5 | GMDN code
6 | GMDN category
7 | UMDNS name
8 | UMDNS code
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Colpomicroscopes, vaginoscopes
11 | Alternative code/s (optional) | N.A.
12 | Keywords (optional) | surgical, gynaecology, magnification
13 | GMDN/UMDNS definition (optional)
Purpose of use
14 | Clinical or other purpose | A colposcope is a low magnification, light-illuminated visualization device for examining the cervix, across an area measuring approximately 20 to 30 mm in diameter, with enough distance between the colposcope lens and the cervix to accommodate the surgical instruments needed for the examination and/or treatment. It allows the examiner to view the epithelial tissues of the cervix and other anogenital areas.
15 | Level of use (if relevant) | Clinic, health centre, (district) hospital, specialized clinic
16 | Clinical department/ ward(if relevant) | Family medicine; gynaecology; outpatient clinic; outreach; oncology; obstetrics; surgery; nursing services
17 | Overview of functional requirements | Colposcope functional requirements: 1) the colposcope head housing the optics capable of magnification and focus 2) the light source and green light filters for improved visualization of vasculature 3) the body or stand to facilitate and optimises biopsy and excisional treatment.
Technical characteristics
18 | Detailed requirements | • Magnification: A range of optical magnification between 3x to 15x (either stepped or continuously variable) • Illumination: o Light sources shall be either halogen or LED to guarantee full-spectrum visible light (white light) - Halogen light: 15 V/150 W - LED: 20,000-35,000 LUX (at 300 mm working distance) o An illumination adjustment knob to change the intensity of light o A fan to cool the bulbs (if halogen bulbs are used) • Green light filters; however, blue filters are also acceptable but not the preferred option • Ingress protection rating: IPX2 (minimum)
19 | Displayed parameters | N.A.
20 | User adjustable settings | • Coarse and fine optical magnification • Illumination
Physical/chemical characteristics
21 | Components(if relevant) | N.A.
22 | Mobility, portability(if relevant) | N.A.
23 | Raw Materials(if relevant) | N.A.
Utility requirements
24 | Electrical, water and/or gas supply (if relevant) | • The unit is suggested to be connected to a reliable power source. • Electrical source requirements (based on country/setting of use): o Amperage: ______ o Voltage: ________ o Plug type: _______
Accessories, consumables, spare parts, other components
25 | Accessories (if relevant) | • Stand or mount to allow for hands-free operation • LED TV or medical grade monitor if not integrated (optional) • Single-use sheath (if using invasive portable model)
26 | Sterilization process for accessories (if relevant) | N.A.
27 | Consumables / reagents (if relevant) | N.A.
28 | Spare parts (if relevant) | Lamp bulbs and fuses.
29 | Other components (if relevant) | N.A.
Packaging
30 | Sterility status on delivery (if relevant) | N.A.
31 | Shelf life (if relevant) | N.A.
32 | Transportation and storage (if relevant) | Storage area should be clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 85% (non-condensing), and be protected from dripping water (IPX2).
33 | Labelling (if relevant) | Labelling on the primary packaging should include the name and/or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements.
Environmental requirements
34 | Context-dependent requirements | Environmental conditions vary globally and can be extreme; however, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
Training, installation and utilization
35 | Pre-installation requirements(if relevant) | N.A.
36 | Requirements for commissioning (if relevant) | N.A.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Clinical staff training in cervical precancer lesion treatment guidelines and device use to be provided.
38 | User care(if relevant) | N.A.
Warranty and maintenance
39 | Warranty | Minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided.
40 | Maintenance tasks | Limited maintenance requirements. Standard colposcopes may require replacement of worn parts including lamps, eyepiece rings, light guides and fuses. Follow device-specific service manual, as instructions are specific to each colposcope model.
41 | Type of service contract | N.A.
42 | Spare parts availability post-warranty | 8 years minimum, starting from installation/commissioning
43 | Software / Hardware upgrade availability | N.A.
Documentation
44 | Documentation requirements | • Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided (if applicable). • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers.
Decommissioning
45 | Estimated Life Span | 10 years
Safety and standards
46 | Standards, for the manufacturer and the equipment | US FDA: Device Class 2 EU: Class Iia
47 | Regulatory Approval / Certification | Compliance to (where applicable, but not limited to): - National Regulatory Agency/Authority (NRA) requirements compliance - Approval by regulatory body of country of manufacturer (if applicable) And at least one of: - FDA 510k clearance (US FDA) - CE mark (EU), with indication of Notifying Body (when applicable) - Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan.
48 | International standards (for reference) | Compliant with active version of the following standards (or equivalent): General manufacturing: • ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes • ISO 14971: Medical Devices - Application of Risk Management to Medical Devices • ISO 15223-1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements Safety & product standards: • IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance • IEC 60601-1-2: Medical electrical equipment - Part 1-2 General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests For vaginally-inserted colposcopes: • Biocompatibility: o ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process o ISO 10993-5: Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity o ISO 10993-10: Biological evaluation of medical devices Part 10: Tests for irritation and skin sensitization • Endoscopy (pertaining to inserted scope): o ISO 8600-1: Endoscopes – Medical endoscopes and endotherapy devices – Part 1: General requirements o ISO 8600-3: Optics and optical instruments – Medical endoscopes and endoscopic accessories – Part 3: Determination of field of view and direction of view of endoscopes with optics o ISO 8600-4: Endoscopes – Medical endoscopes and endotherapy devices – Part 4: Determination of maximum width of insertion portion o ISO 8600-5: Optics and photonics – Medical endoscopes and endotherapy devices – Part 5: Determination of optical resolution of rigid endoscopes with optics o ISO 8600-6: Optics and photonics – Medical endoscopes and endotherapy devices – Part 6: vocabulary
49 | Regional/local standards | Country-specific and regional standards may apply
50 | Regulations | US regulations: 21 CFR part 820 Quality System Regulation 21 CFR part 884.1630 - Colposcope EU regulations: European Commission Regulation (EU) No. 2017/745 (replacing 93/42/EEC)